## Comments & Info
Measurement information
Contacts :
SIMS Tool: | Cameca | Author: | B.Douhard
Primary Beam: | Cs+ | Reviewed by:
Mass resolution | 4000 HMR
Extra | Electron Gun
Comments:
General remarks: | Please note that plots show in the same time quantified (As From 75As- or 103AsSi-, At./cm3) and unquantified curves (Counts/s).
Samples As-1 and As-2 were measured twice in order to asses reproducibility and to determine wich concentration quantification is the most accurate. All measurements are shown separately under red tabs. It should be noted that these were very challenging measurements: Low primary energy for optimal depth resolution + High mass resolution requiring very narrow energy band-pass making signal very sensitive to surface charging. We used electron gun to compensate for charging.
Depthscales: | Concerning depth, it must be kept in mind that absolute values are not accurate (a few nm precision because of Sputter Rate (SR, nm/min) variation during transiants) but relative depths between samples are accurate.. For this request, Standard implant (plotted under blue tab)depthscale was determined using total crater measured depth (Dektak profilometer) and assuming constant sputter rate along profile. Samples measurements were not calibrated in depth since scale would be distorted becasue of varying sputter rates through different layers. We still measured their total dephts as an indication (with rather large error bars because of their low values and poor reproducibility (Roughness, Gold layer scratching with stylus...). Measured depths and SR values can be found under the "Depths" tab.
Concentration quantification: | Arsenic was measured two ways: Following either 75As- or 103AsSi- ions. Both were quantified to represent As concentration using our reference As in Si implant. Note that both methods should match when As is diluted in Si matrix but AsSi can give an underestimate if we have As clusters for instance. Also, these two ions can be quantified using simple constant conversion from Counts/s to At./cm3 (Sensitivity Factor method, SF) or including extra normalization relative to Si reference signal (Relative Sensitivity Factor method, RSF). We tried both SF and RSF methods and used repeated measurements of As-1 and As-2 samples to see which method gave better As concentration reproducibility (ie if we had to correct or not As and AsSi for Si level variations between measurements). Based on As-1 and As-1 bis measurements (see 1st grey tab), it appears SF method is more reproducible. We then used SF method to compare samples (first two black tabs). In case comparison using RSF are also visible under light grey tabs before datasheets. Concerning As vs AsSi we can see that both methods match on right hand side of peak (ie where As is diluted in Si matrix) as they should while onset of single As arrives sooner than the one of AsSi. This can be linked with clustered As giving As signal but no significant AsSi signal. Any quantified signal is only meaningful where we can assume same matrix as for the standard sample used for quantification (Here only if we can assume As diluted in Si what may not be true at implant peak). At top surface and interfaces, SIMS artefacts are present that make the quantification deviating from real level.
Interpretation | Global conclusion is that we could no evidence clear differnces between samples (measured differences between samples being of the same order as differences observed between 2 runs on same sample.
Feel free to set a meeting..
General info about SIMS at Imec: | Info SIMS Imec

## Depths
| Sample | AsinSiO2 Ref 5e14 | As-1 |  | As-2 |  | As-3
| Measurement | 14 | 7 | 11 | 9 | 12 | 10
Dektak Depths | Avge Depth (nm) | 103.69 | 26.28 | 21.799 | 18.05 | 16.266 | 26.3 | Depths values in red have unusual large error bars because of their low values and poor reproducibility (Roughness, Gold layer scratching with stylus...)
| Tot Meas.Time (min) | 120.225
Sputter Rate Check | Sputter Rate (nm/min) | 0.86247
Primary Current
| SR/Ip | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Quantification | Matrix Signal Intensity | 5278
| Rel.Sens.Factor AsSi | 6.926e+21
| Rel.Sens.Factor As | 2.559e+21
7213.5
120.225
Please refer to "Comments & Info" tab

## Data_AsinSiO2
*** DATA FILES ***
Raw data file name(s) | 14_AsinSiO2_NEG GOOD.ckb
Sample ID |  | SIC13_047_3-1_Al2O3
Sample code
Data process ID
Analysis date |  | 2017-09-14
Analysis time |  | 16:14
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
14_AsinSiO2_NEG GOOD.ckb
AsinSiO2 12C Int. (Counts/s) |  |  | AsinSiO2 13C Int. (Counts/s) |  |  | AsinSiO2 16O Int. (Counts/s) |  |  | AsinSiO2 18O Int. (Counts/s) |  |  | AsinSiO2 28Si Int. (Counts/s) |  |  | AsinSiO2 75As Conc. (At./cm3) |  |  | AsinSiO2 As Conc. from 103AsSi (At./cm3) |  |  | AsinSiO2 197Au Int. (Counts/s) |  |  |  |  | As-2 75As 28Si_RSF Int. (Counts/s)
Time | Depth[nm] | I[cnt/s] | Time | Depth[nm] | I[cnt/s] | Time | Depth[nm] | I[cnt/s] | Time | Depth[nm] | I[cnt/s] | Time | Depth[nm] | I[cnt/s] | Time | Depth[nm] | C[atom/cm3] | Time | Depth[nm] | C[atom/cm3] | Time | Depth[nm] | I[cnt/s]
| 0.11614 | 4.2604 |  | 0.21044 | 0.20288 |  | 0.26219 | 16.3604 |  | 0.35763 | 0 |  | 0.39673 | 6.8826 |  | 0.49678 | 0 |  | 0.59337 | 2.6624e+17 |  | 0.66582 | 0
| 0.78082 | 8.7238 |  | 0.87511 | 2.2317 |  | 0.92686 | 5209.29 |  | 1.0223 | 13.7957 |  | 1.0614 | 302.841 |  | 1.1615 | 98378600000000000 |  | 1.258 | 7.9871e+17 |  | 1.3305 | 0
| 1.4455 | 2.8403 |  | 1.5398 | 1.623 |  | 1.5915 | 20315.5 |  | 1.687 | 46.2564 |  | 1.7261 | 545.709 |  | 1.8261 | 1.9676e+17 |  | 1.9227 | 5.3247e+17 |  | 1.9952 | 0
| 2.1102 | 2.8403 |  | 2.2045 | 2.2317 |  | 2.2562 | 31502.1 |  | 2.3516 | 63.9069 |  | 2.3908 | 669.604 |  | 2.4908 | 0 |  | 2.5874 | 2.6624e+17 |  | 2.6598 | 0
| 2.7748 | 3.0432 |  | 2.8691 | 3.0432 |  | 2.9209 | 38925.3 |  | 3.0163 | 80.3402 |  | 3.0554 | 818.083 |  | 3.1555 | 0 |  | 3.2521 | 2.6624e+17 |  | 3.3245 | 0
| 3.4395 | 4.4633 |  | 3.5338 | 2.0288 |  | 3.5856 | 42981.7 |  | 3.681 | 87.847 |  | 3.7201 | 890.845 |  | 3.8201 | 2.9514e+17 |  | 3.9167 | 0 |  | 3.9892 | 0
| 4.1042 | 3.6518 |  | 4.1985 | 2.6374 |  | 4.2502 | 49163.4 |  | 4.3457 | 90.6871 |  | 4.3848 | 1054.08 |  | 4.4848 | 0 |  | 4.5814 | 7.9871e+17 |  | 4.6539 | 0
| 4.7688 | 4.8691 |  | 4.8632 | 3.0432 |  | 4.9149 | 61133.1 |  | 5.0103 | 121.525 |  | 5.0494 | 1285.17 |  | 5.1495 | 2.9514e+17 |  | 5.2461 | 0 |  | 5.3185 | 0
| 5.4335 | 3.8547 |  | 5.5278 | 5.072 |  | 5.5796 | 64693.6 |  | 5.675 | 120.105 |  | 5.7141 | 1320.56 |  | 5.8142 | 98378600000000000 |  | 5.9108 | 0 |  | 5.9832 | 0
| 6.0982 | 7.5065 |  | 6.1925 | 5.2748 |  | 6.2442 | 69822.6 |  | 6.3397 | 138.364 |  | 6.3788 | 1426.76 |  | 6.4788 | 1.9676e+17 |  | 6.5754 | 0 |  | 6.6479 | 0
| 6.7629 | 5.8835 |  | 6.8572 | 4.8691 |  | 6.9089 | 76100.6 |  | 7.0044 | 157.638 |  | 7.0434 | 1591.98 |  | 7.1435 | 1.9676e+17 |  | 7.2401 | 7.9871e+17 |  | 7.3125 | 0
| 7.4275 | 6.2892 |  | 7.5218 | 5.6806 |  | 7.5736 | 86280.7 |  | 7.669 | 174.072 |  | 7.7081 | 1804.38 |  | 7.8082 | 2.9514e+17 |  | 7.9048 | 5.3247e+17 |  | 7.9772 | 0
| 8.0922 | 6.2892 |  | 8.1865 | 6.4921 |  | 8.2383 | 95756.5 |  | 8.3337 | 192.331 |  | 8.3728 | 2028.57 |  | 8.4728 | 8.8541e+17 |  | 8.5694 | 5.3247e+17 |  | 8.6419 | 0
| 8.7569 | 7.3036 |  | 8.8512 | 6.2892 |  | 8.9029 | 101765 |  | 8.9984 | 198.214 |  | 9.0375 | 3008.96 |  | 9.1375 | 8.8541e+17 |  | 9.2341 | 2.6624e+17 |  | 9.3066 | 0
| 9.4216 | 11.1583 |  | 9.5159 | 11.7669 |  | 9.5676 | 160101 |  | 9.663 | 303.103 |  | 9.7021 | 3110.24 |  | 9.8022 | 1.6724e+18 |  | 9.8988 | 1.0649e+18 |  | 9.9712 | 0
| 10.0862 | 9.9411 |  | 10.1805 | 10.5497 |  | 10.2323 | 166128 |  | 10.3277 | 338.81 |  | 10.3668 | 3180.06 |  | 10.4669 | 1.9676e+18 |  | 10.5635 | 2.1299e+18 |  | 10.6359 | 0
| 10.7509 | 11.9698 |  | 10.8452 | 8.9267 |  | 10.8969 | 167529 |  | 10.9924 | 331.101 |  | 11.0315 | 3305.93 |  | 11.1315 | 1.8692e+18 |  | 11.2281 | 1.5974e+18 |  | 11.3006 | 0
| 11.4156 | 15.4188 |  | 11.5099 | 9.941 |  | 11.5616 | 183115 |  | 11.6571 | 379.793 |  | 11.6962 | 3500.65 |  | 11.7962 | 2.066e+18 |  | 11.8928 | 2.1299e+18 |  | 11.9652 | 0
| 12.0802 | 14.4044 |  | 12.1745 | 8.7238 |  | 12.2263 | 183188 |  | 12.3217 | 373.098 |  | 12.3608 | 3492.78 |  | 12.4609 | 3.5416e+18 |  | 12.5575 | 1.8636e+18 |  | 12.6299 | 0
| 12.7449 | 13.9986 |  | 12.8392 | 8.318 |  | 12.891 | 191827 |  | 12.9864 | 389.328 |  | 13.0255 | 3829.12 |  | 13.1255 | 2.6562e+18 |  | 13.2221 | 3.7273e+18 |  | 13.2946 | 0
| 13.4096 | 10.9554 |  | 13.5039 | 10.3468 |  | 13.5556 | 197015 |  | 13.6511 | 398.255 |  | 13.6902 | 3795.68 |  | 13.7902 | 3.7384e+18 |  | 13.8868 | 2.3961e+18 |  | 13.9593 | 0
| 14.0743 | 14.2015 |  | 14.1686 | 8.7238 |  | 14.2203 | 201324 |  | 14.3157 | 407.182 |  | 14.3548 | 3981.55 |  | 14.4549 | 4.8206e+18 |  | 14.5515 | 5.0585e+18 |  | 14.6239 | 0
| 14.7389 | 11.767 |  | 14.8332 | 8.9267 |  | 14.885 | 199201 |  | 14.9804 | 389.328 |  | 15.0195 | 3941.23 |  | 15.1196 | 5.706e+18 |  | 15.2162 | 4.7922e+18 |  | 15.2886 | 0
| 15.4036 | 11.9698 |  | 15.4979 | 11.1583 |  | 15.5496 | 209914 |  | 15.6451 | 430.919 |  | 15.6842 | 4071.02 |  | 15.7842 | 5.8043e+18 |  | 15.8808 | 5.591e+18 |  | 15.9533 | 0
| 16.0683 | 12.1727 |  | 16.1626 | 15.013 |  | 16.2143 | 208994 |  | 16.3098 | 406.167 |  | 16.3489 | 4070.06 |  | 16.4489 | 6.7881e+18 |  | 16.5455 | 7.7208e+18 |  | 16.618 | 0
| 16.7329 | 14.4044 |  | 16.8272 | 9.7382 |  | 16.879 | 208834 |  | 16.9744 | 424.63 |  | 17.0135 | 4068.06 |  | 17.1136 | 7.8703e+18 |  | 17.2102 | 1.2247e+19 |  | 17.2826 | 0
| 17.3976 | 16.4332 |  | 17.4919 | 13.1871 |  | 17.5437 | 211012 |  | 17.6391 | 437.209 |  | 17.6782 | 4195.94 |  | 17.7783 | 9.8379e+18 |  | 17.8748 | 1.0117e+19 |  | 17.9473 | 0
| 18.0623 | 16.0274 |  | 18.1566 | 10.9554 |  | 18.2083 | 214908 |  | 18.3038 | 440.658 |  | 18.3429 | 4337.55 |  | 18.4429 | 1.1904e+19 |  | 18.5395 | 1.4909e+19 |  | 18.612 | 0
| 18.727 | 14.2015 |  | 18.8213 | 15.8245 |  | 18.873 | 219763 |  | 18.9685 | 445.121 |  | 19.0076 | 4305.1 |  | 19.1076 | 1.151e+19 |  | 19.2042 | 1.3312e+19 |  | 19.2766 | 0
| 19.3916 | 15.4188 |  | 19.4859 | 11.9698 |  | 19.5377 | 224357 |  | 19.6331 | 440.049 |  | 19.6722 | 4293.3 |  | 19.7723 | 1.3773e+19 |  | 19.8689 | 1.5974e+19 |  | 19.9413 | 0
| 20.0563 | 18.2591 |  | 20.1506 | 15.8245 |  | 20.2024 | 225509 |  | 20.2978 | 485.292 |  | 20.3369 | 4333.6 |  | 20.4369 | 1.761e+19 |  | 20.5335 | 1.8903e+19 |  | 20.606 | 0
| 20.721 | 17.0418 |  | 20.8153 | 15.013 |  | 20.867 | 228987 |  | 20.9625 | 470.076 |  | 21.0016 | 4748.65 |  | 21.1016 | 1.879e+19 |  | 21.1982 | 1.9701e+19 |  | 21.2707 | 0
| 21.3857 | 15.4188 |  | 21.4799 | 12.1727 |  | 21.5317 | 238624 |  | 21.6271 | 476.771 |  | 21.6662 | 4564.74 |  | 21.7663 | 1.8889e+19 |  | 21.8629 | 2.3429e+19 |  | 21.9353 | 0
| 22.0503 | 15.6216 |  | 22.1446 | 13.5929 |  | 22.1964 | 232795 |  | 22.2918 | 481.235 |  | 22.3309 | 4602.12 |  | 22.431 | 1.8003e+19 |  | 22.5276 | 2.689e+19 |  | 22.6 | 0
| 22.715 | 13.1871 |  | 22.8093 | 13.39 |  | 22.861 | 233298 |  | 22.9565 | 462.975 |  | 22.9956 | 4596.21 |  | 23.0956 | 2.2824e+19 |  | 23.1922 | 3.2214e+19 |  | 23.2647 | 0
| 23.3797 | 12.3756 |  | 23.474 | 14.6073 |  | 23.5257 | 236294 |  | 23.6212 | 484.481 |  | 23.6603 | 4757.49 |  | 23.7603 | 2.7152e+19 |  | 23.8569 | 3.115e+19 |  | 23.9293 | 0
| 24.0443 | 14.2015 |  | 24.1386 | 16.0274 |  | 24.1904 | 237616 |  | 24.2858 | 503.957 |  | 24.3249 | 4708.34 |  | 24.425 | 3.0006e+19 |  | 24.5216 | 3.115e+19 |  | 24.594 | 0
| 24.709 | 12.9842 |  | 24.8033 | 10.9554 |  | 24.8551 | 243627 |  | 24.9505 | 513.696 |  | 24.9896 | 4855.82 |  | 25.0896 | 3.3941e+19 |  | 25.1862 | 3.2215e+19 |  | 25.2587 | 0
| 25.3737 | 17.0418 |  | 25.468 | 15.6216 |  | 25.5197 | 245444 |  | 25.6152 | 494.828 |  | 25.6543 | 4868.64 |  | 25.7543 | 3.3252e+19 |  | 25.8509 | 3.7007e+19 |  | 25.9234 | 0
| 26.0384 | 16.636 |  | 26.1327 | 14.8101 |  | 26.1844 | 248095 |  | 26.2798 | 501.726 |  | 26.3189 | 4966.98 |  | 26.419 | 3.9155e+19 |  | 26.5156 | 4.1e+19 |  | 26.588 | 0
| 26.703 | 17.8533 |  | 26.7973 | 15.2159 |  | 26.8491 | 251348 |  | 26.9445 | 498.48 |  | 26.9836 | 4977.8 |  | 27.0837 | 4.1024e+19 |  | 27.1803 | 4.4728e+19 |  | 27.2527 | 0
| 27.3677 | 19.6792 |  | 27.462 | 10.1439 |  | 27.5137 | 252876 |  | 27.6092 | 515.928 |  | 27.6483 | 5037.79 |  | 27.7483 | 4.9583e+19 |  | 27.8449 | 4.9786e+19 |  | 27.9174 | 0
| 28.0324 | 16.0274 |  | 28.1267 | 15.2159 |  | 28.1784 | 248608 |  | 28.2739 | 515.522 |  | 28.313 | 5138.11 |  | 28.413 | 4.7419e+19 |  | 28.5096 | 5.165e+19 |  | 28.5821 | 0
| 28.697 | 18.6648 |  | 28.7913 | 10.9554 |  | 28.8431 | 253804 |  | 28.9385 | 524.246 |  | 28.9776 | 4987.63 |  | 29.0777 | 5.1452e+19 |  | 29.1743 | 5.7773e+19 |  | 29.2467 | 0
| 29.3617 | 21.5051 |  | 29.456 | 13.9986 |  | 29.5078 | 258176 |  | 29.6032 | 525.463 |  | 29.6423 | 5095.82 |  | 29.7424 | 5.7945e+19 |  | 29.8389 | 5.9903e+19 |  | 29.9114 | 0
| 30.0264 | 15.4188 |  | 30.1207 | 18.0562 |  | 30.1724 | 257144 |  | 30.2679 | 545.143 |  | 30.307 | 5268.91 |  | 30.407 | 6.2667e+19 |  | 30.5036 | 6.1234e+19 |  | 30.5761 | 0
| 30.6911 | 17.8533 |  | 30.7854 | 13.1871 |  | 30.8371 | 260957 |  | 30.9326 | 536.216 |  | 30.9716 | 5032.88 |  | 31.0717 | 6.434e+19 |  | 31.1683 | 6.7358e+19 |  | 31.2407 | 0
| 31.3557 | 18.8677 |  | 31.45 | 17.0418 |  | 31.5018 | 258130 |  | 31.5972 | 530.738 |  | 31.6363 | 5197.11 |  | 31.7364 | 6.7783e+19 |  | 31.833 | 7.5877e+19 |  | 31.9054 | 0
| 32.0204 | 16.8389 |  | 32.1147 | 16.8389 |  | 32.1665 | 258526 |  | 32.2619 | 544.129 |  | 32.301 | 5199.09 |  | 32.401 | 7.1718e+19 |  | 32.4976 | 6.8955e+19 |  | 32.5701 | 0
| 32.6851 | 17.2447 |  | 32.7794 | 14.8101 |  | 32.8311 | 265586 |  | 32.9266 | 533.984 |  | 32.9657 | 5271.85 |  | 33.0657 | 7.4473e+19 |  | 33.1623 | 7.8274e+19 |  | 33.2348 | 0
| 33.3498 | 19.6792 |  | 33.444 | 15.013 |  | 33.4958 | 265531 |  | 33.5912 | 549.2 |  | 33.6303 | 5462.66 |  | 33.7304 | 8.3524e+19 |  | 33.827 | 8.2267e+19 |  | 33.8994 | 0
| 34.0144 | 18.8677 |  | 34.1087 | 13.1871 |  | 34.1605 | 266115 |  | 34.2559 | 566.649 |  | 34.295 | 5506.94 |  | 34.3951 | 7.9982e+19 |  | 34.4917 | 7.428e+19 |  | 34.5641 | 0
| 34.6791 | 19.4763 |  | 34.7734 | 13.1871 |  | 34.8251 | 264730 |  | 34.9206 | 556.707 |  | 34.9597 | 5428.25 |  | 35.0597 | 8.4508e+19 |  | 35.1563 | 9.7176e+19 |  | 35.2288 | 0
| 35.3438 | 17.4476 |  | 35.4381 | 15.2159 |  | 35.4898 | 264153 |  | 35.5853 | 552.244 |  | 35.6244 | 5417.45 |  | 35.7244 | 8.2835e+19 |  | 35.821 | 8.4663e+19 |  | 35.8934 | 0
| 36.0084 | 18.4619 |  | 36.1027 | 17.6504 |  | 36.1545 | 270725 |  | 36.2499 | 555.287 |  | 36.289 | 5415.46 |  | 36.3891 | 9.4542e+19 |  | 36.4857 | 9.1852e+19 |  | 36.5581 | 0
| 36.6731 | 18.4619 |  | 36.7674 | 15.4188 |  | 36.8192 | 272830 |  | 36.9146 | 568.069 |  | 36.9537 | 5608.24 |  | 37.0537 | 9.6412e+19 |  | 37.1503 | 9.7975e+19 |  | 37.2228 | 0
| 37.3378 | 19.6792 |  | 37.4321 | 15.6216 |  | 37.4838 | 271159 |  | 37.5793 | 559.751 |  | 37.6184 | 5580.69 |  | 37.7184 | 9.346e+19 |  | 37.815 | 8.4663e+19 |  | 37.8875 | 0
| 38.0025 | 19.2735 |  | 38.0968 | 16.2303 |  | 38.1485 | 270705 |  | 38.2439 | 556.302 |  | 38.283 | 5549.22 |  | 38.3831 | 1.0625e+20 |  | 38.4797 | 9.2916e+19 |  | 38.5521 | 0
| 38.6671 | 20.6936 |  | 38.7614 | 17.0418 |  | 38.8132 | 268085 |  | 38.9086 | 526.883 |  | 38.9477 | 5308.26 |  | 39.0478 | 9.6707e+19 |  | 39.1444 | 9.8508e+19 |  | 39.2168 | 0
| 39.3318 | 17.4476 |  | 39.4261 | 19.2735 |  | 39.4778 | 270119 |  | 39.5733 | 559.547 |  | 39.6124 | 5388.89 |  | 39.7124 | 1.0241e+20 |  | 39.809 | 9.7176e+19 |  | 39.8815 | 0
| 39.9965 | 19.2735 |  | 40.0908 | 15.8245 |  | 40.1425 | 268046 |  | 40.238 | 534.39 |  | 40.2771 | 5334.79 |  | 40.3771 | 1.0477e+20 |  | 40.4737 | 1.1022e+20 |  | 40.5462 | 0
| 40.6611 | 19.8821 |  | 40.7554 | 13.39 |  | 40.8072 | 265810 |  | 40.9026 | 537.028 |  | 40.9417 | 5397.74 |  | 41.0418 | 1.0527e+20 |  | 41.1384 | 9.9306e+19 |  | 41.2108 | 0
| 41.3258 | 21.9109 |  | 41.4201 | 16.0274 |  | 41.4719 | 264520 |  | 41.5673 | 548.389 |  | 41.6064 | 5375.13 |  | 41.7064 | 1.1078e+20 |  | 41.803 | 9.1319e+19 |  | 41.8755 | 0
| 41.9905 | 15.8245 |  | 42.0848 | 14.6072 |  | 42.1365 | 264416 |  | 42.232 | 540.882 |  | 42.2711 | 5246.3 |  | 42.3711 | 1.0881e+20 |  | 42.4677 | 1.1634e+20 |  | 42.5402 | 0
| 42.6552 | 21.5051 |  | 42.7495 | 16.4332 |  | 42.8012 | 262043 |  | 42.8967 | 546.969 |  | 42.9357 | 5501.05 |  | 43.0358 | 1.1107e+20 |  | 43.1324 | 1.0889e+20 |  | 43.2048 | 0
| 43.3198 | 18.8677 |  | 43.4141 | 15.4188 |  | 43.4659 | 264419 |  | 43.5613 | 559.345 |  | 43.6004 | 5545.28 |  | 43.7005 | 1.1284e+20 |  | 43.7971 | 1.1475e+20 |  | 43.8695 | 0
| 43.9845 | 19.2735 |  | 44.0788 | 14.4044 |  | 44.1306 | 271678 |  | 44.226 | 550.621 |  | 44.2651 | 5464.65 |  | 44.3651 | 1.1127e+20 |  | 44.4617 | 1.1448e+20 |  | 44.5342 | 0
| 44.6492 | 19.2735 |  | 44.7435 | 16.4332 |  | 44.7952 | 270265 |  | 44.8907 | 554.273 |  | 44.9298 | 5533.5 |  | 45.0298 | 1.0999e+20 |  | 45.1264 | 1.1075e+20 |  | 45.1989 | 0
| 45.3138 | 19.0706 |  | 45.4081 | 15.2159 |  | 45.4599 | 266494 |  | 45.5553 | 529.318 |  | 45.5944 | 5379.04 |  | 45.6945 | 1.1333e+20 |  | 45.7911 | 1.1448e+20 |  | 45.8635 | 0
| 45.9785 | 15.6216 |  | 46.0728 | 13.1871 |  | 46.1246 | 271281 |  | 46.22 | 540.071 |  | 46.2591 | 5489.24 |  | 46.3592 | 1.216e+20 |  | 46.4558 | 1.1049e+20 |  | 46.5282 | 0
| 46.6432 | 15.2159 |  | 46.7375 | 16.0274 |  | 46.7892 | 267679 |  | 46.8847 | 554.881 |  | 46.9238 | 5412.51 |  | 47.0238 | 1.1166e+20 |  | 47.1204 | 1.0996e+20 |  | 47.1929 | 0
| 47.3079 | 16.8389 |  | 47.4022 | 16.8389 |  | 47.4539 | 267165 |  | 47.5494 | 551.839 |  | 47.5885 | 5315.14 |  | 47.6885 | 1.1422e+20 |  | 47.7851 | 1.1821e+20 |  | 47.8575 | 0
| 47.9725 | 15.013 |  | 48.0668 | 16.2303 |  | 48.1186 | 269222 |  | 48.214 | 553.258 |  | 48.2531 | 5295.46 |  | 48.3532 | 1.2052e+20 |  | 48.4498 | 1.1208e+20 |  | 48.5222 | 0
| 48.6372 | 15.8245 |  | 48.7315 | 12.7813 |  | 48.7833 | 265635 |  | 48.8787 | 536.825 |  | 48.9178 | 5471.52 |  | 49.0178 | 1.1156e+20 |  | 49.1144 | 1.0703e+20 |  | 49.1869 | 0
| 49.3019 | 19.2735 |  | 49.3962 | 15.6217 |  | 49.4479 | 266947 |  | 49.5434 | 532.158 |  | 49.5825 | 5348.58 |  | 49.6825 | 1.1314e+20 |  | 49.7791 | 1.0889e+20 |  | 49.8516 | 0
| 49.9666 | 18.8677 |  | 50.0608 | 13.9986 |  | 50.1126 | 267513 |  | 50.208 | 548.186 |  | 50.2471 | 5342.66 |  | 50.3472 | 1.1333e+20 |  | 50.4438 | 1.1661e+20 |  | 50.5162 | 0
| 50.6312 | 14.6072 |  | 50.7255 | 13.5929 |  | 50.7773 | 269193 |  | 50.8727 | 547.578 |  | 50.9118 | 5247.29 |  | 51.0119 | 1.1412e+20 |  | 51.1085 | 1.0836e+20 |  | 51.1809 | 0
| 51.2959 | 14.8101 |  | 51.3902 | 17.6504 |  | 51.4419 | 265535 |  | 51.5374 | 574.155 |  | 51.5765 | 5452.85 |  | 51.6765 | 1.0979e+20 |  | 51.7731 | 1.1288e+20 |  | 51.8456 | 0
| 51.9606 | 16.2303 |  | 52.0549 | 16.0274 |  | 52.1066 | 268776 |  | 52.2021 | 509.029 |  | 52.2412 | 5364.32 |  | 52.3412 | 1.1402e+20 |  | 52.4378 | 1.1634e+20 |  | 52.5102 | 0
| 52.6252 | 21.708 |  | 52.7195 | 15.6216 |  | 52.7713 | 274041 |  | 52.8667 | 548.998 |  | 52.9058 | 5263 |  | 53.0059 | 1.1156e+20 |  | 53.1025 | 1.0756e+20 |  | 53.1749 | 0
| 53.2899 | 19.0706 |  | 53.3842 | 16.8389 |  | 53.436 | 275083 |  | 53.5314 | 548.592 |  | 53.5705 | 5498.09 |  | 53.6705 | 1.1048e+20 |  | 53.7671 | 1.0889e+20 |  | 53.8396 | 0
| 53.9546 | 19.6792 |  | 54.0489 | 15.013 |  | 54.1006 | 275889 |  | 54.1961 | 527.086 |  | 54.2352 | 5231.55 |  | 54.3352 | 1.0999e+20 |  | 54.4318 | 1.1395e+20 |  | 54.5043 | 0
| 54.6193 | 17.6504 |  | 54.7136 | 14.6073 |  | 54.7653 | 268960 |  | 54.8607 | 560.156 |  | 54.8998 | 5282.67 |  | 54.9999 | 1.0782e+20 |  | 55.0965 | 1.0916e+20 |  | 55.1689 | 0
| 55.2839 | 16.0274 |  | 55.3782 | 12.9842 |  | 55.43 | 269885 |  | 55.5254 | 542.708 |  | 55.5645 | 5236.46 |  | 55.6646 | 1.0399e+20 |  | 55.7612 | 1.0756e+20 |  | 55.8336 | 0
| 55.9486 | 18.0562 |  | 56.0429 | 12.7813 |  | 56.0946 | 269397 |  | 56.1901 | 542.1 |  | 56.2292 | 5359.4 |  | 56.3292 | 1.0448e+20 |  | 56.4258 | 9.8508e+19 |  | 56.4983 | 0
| 56.6133 | 13.9986 |  | 56.7076 | 15.6216 |  | 56.7593 | 268893 |  | 56.8548 | 549.404 |  | 56.8939 | 5309.27 |  | 56.9939 | 1.0832e+20 |  | 57.0905 | 1.0889e+20 |  | 57.163 | 0
| 57.2779 | 16.0274 |  | 57.3722 | 13.39 |  | 57.424 | 262568 |  | 57.5194 | 537.434 |  | 57.5585 | 5162.69 |  | 57.6586 | 1.0074e+20 |  | 57.7552 | 1.0596e+20 |  | 57.8276 | 0
| 57.9426 | 16.636 |  | 58.0369 | 15.2159 |  | 58.0887 | 262814 |  | 58.1841 | 539.259 |  | 58.2232 | 5238.42 |  | 58.3232 | 1.0182e+20 |  | 58.4198 | 9.5313e+19 |  | 58.4923 | 0
| 58.6073 | 18.0562 |  | 58.7016 | 15.8245 |  | 58.7533 | 267629 |  | 58.8488 | 560.157 |  | 58.8879 | 5271.88 |  | 58.9879 | 1.0094e+20 |  | 59.0845 | 9.4514e+19 |  | 59.157 | 0
| 59.272 | 20.2878 |  | 59.3663 | 16.8389 |  | 59.418 | 267467 |  | 59.5135 | 530.738 |  | 59.5526 | 5121.4 |  | 59.6526 | 9.3362e+19 |  | 59.7492 | 9.1585e+19 |  | 59.8216 | 0
| 59.9366 | 18.462 |  | 60.0309 | 16.636 |  | 60.0827 | 264223 |  | 60.1781 | 546.157 |  | 60.2172 | 5267.93 |  | 60.3173 | 8.5983e+19 |  | 60.4139 | 8.8124e+19 |  | 60.4863 | 0
| 60.6013 | 17.0418 |  | 60.6956 | 13.39 |  | 60.7474 | 266405 |  | 60.8428 | 545.752 |  | 60.8819 | 5080.07 |  | 60.9819 | 9.651e+19 |  | 61.0785 | 9.1053e+19 |  | 61.151 | 0
| 61.266 | 15.6216 |  | 61.3603 | 12.1727 |  | 61.412 | 269871 |  | 61.5075 | 541.288 |  | 61.5466 | 5108.6 |  | 61.6466 | 9.2772e+19 |  | 61.7432 | 9.7176e+19 |  | 61.8157 | 0
| 61.9306 | 20.085 |  | 62.0249 | 13.9986 |  | 62.0767 | 267504 |  | 62.1721 | 554.476 |  | 62.2112 | 4990.6 |  | 62.3113 | 8.8836e+19 |  | 62.4079 | 8.9189e+19 |  | 62.4803 | 0
| 62.5953 | 22.3166 |  | 62.6896 | 16.8389 |  | 62.7414 | 265779 |  | 62.8368 | 545.346 |  | 62.8759 | 5214.83 |  | 62.976 | 7.9884e+19 |  | 63.0726 | 7.7208e+19 |  | 63.145 | 0
| 63.26 | 18.4619 |  | 63.3543 | 14.4044 |  | 63.406 | 269259 |  | 63.5015 | 536.013 |  | 63.5406 | 5120.39 |  | 63.6406 | 8.2934e+19 |  | 63.7372 | 7.3481e+19 |  | 63.8097 | 0
| 63.9247 | 20.2878 |  | 64.019 | 17.0418 |  | 64.0707 | 266833 |  | 64.1662 | 539.462 |  | 64.2053 | 4800.77 |  | 64.3053 | 8.0868e+19 |  | 64.4019 | 8.5994e+19 |  | 64.4743 | 0
| 64.5893 | 16.636 |  | 64.6836 | 13.9986 |  | 64.7354 | 260150 |  | 64.8308 | 523.434 |  | 64.8699 | 5065.34 |  | 64.97 | 7.7031e+19 |  | 65.0666 | 7.641e+19 |  | 65.139 | 0
| 65.254 | 19.6792 |  | 65.3483 | 16.636 |  | 65.4001 | 257495 |  | 65.4955 | 527.289 |  | 65.5346 | 4994.51 |  | 65.6346 | 7.3784e+19 |  | 65.7312 | 7.4546e+19 |  | 65.8037 | 0.33635
| 65.9187 | 19.6792 |  | 66.013 | 11.5641 |  | 66.0647 | 262393 |  | 66.1602 | 521.609 |  | 66.1993 | 4995.49 |  | 66.2993 | 7.7424e+19 |  | 66.3959 | 7.4546e+19 |  | 66.4684 | 0
| 66.5834 | 17.8533 |  | 66.6777 | 17.0418 |  | 66.7294 | 258853 |  | 66.8248 | 513.29 |  | 66.8639 | 4956.16 |  | 66.964 | 6.798e+19 |  | 67.0606 | 7.4812e+19 |  | 67.133 | 0
| 67.248 | 20.8965 |  | 67.3423 | 16.0274 |  | 67.3941 | 259445 |  | 67.4895 | 519.782 |  | 67.5286 | 4833.23 |  | 67.6287 | 6.9948e+19 |  | 67.7253 | 6.1234e+19 |  | 67.7977 | 0
| 67.9127 | 16.636 |  | 68.007 | 16.4332 |  | 68.0587 | 264849 |  | 68.1542 | 530.739 |  | 68.1933 | 4921.74 |  | 68.2933 | 6.4241e+19 |  | 68.3899 | 6.5228e+19 |  | 68.4624 | 0
| 68.5774 | 19.2735 |  | 68.6717 | 13.5929 |  | 68.7234 | 264856 |  | 68.8189 | 526.883 |  | 68.858 | 4945.35 |  | 68.958 | 6.9357e+19 |  | 69.0546 | 6.363e+19 |  | 69.1271 | 0
| 69.242 | 17.0418 |  | 69.3363 | 15.2159 |  | 69.3881 | 262761 |  | 69.4835 | 533.173 |  | 69.5226 | 4918.81 |  | 69.6227 | 6.7685e+19 |  | 69.7193 | 6.363e+19 |  | 69.7917 | 0
| 69.9067 | 17.4475 |  | 70.001 | 11.9698 |  | 70.0528 | 260238 |  | 70.1482 | 529.521 |  | 70.1873 | 4821.43 |  | 70.2873 | 6.0208e+19 |  | 70.3839 | 6.2566e+19 |  | 70.4564 | 0
| 70.5714 | 17.6504 |  | 70.6657 | 12.3756 |  | 70.7174 | 263216 |  | 70.8129 | 521 |  | 70.852 | 4955.2 |  | 70.952 | 5.8339e+19 |  | 71.0486 | 5.5643e+19 |  | 71.1211 | 0
| 71.2361 | 18.6648 |  | 71.3304 | 12.7813 |  | 71.3821 | 257869 |  | 71.4776 | 519.783 |  | 71.5166 | 4820.44 |  | 71.6167 | 5.4108e+19 |  | 71.7133 | 4.952e+19 |  | 71.7857 | 0
| 71.9007 | 17.8533 |  | 71.995 | 15.2159 |  | 72.0468 | 251188 |  | 72.1422 | 493.813 |  | 72.1813 | 4730.94 |  | 72.2814 | 5.3321e+19 |  | 72.378 | 4.9254e+19 |  | 72.4504 | 0.33635
| 72.5654 | 17.6504 |  | 72.6597 | 11.9698 |  | 72.7114 | 251988 |  | 72.8069 | 488.133 |  | 72.846 | 4658.15 |  | 72.946 | 5.5879e+19 |  | 73.0426 | 5.9371e+19 |  | 73.1151 | 0
| 73.2301 | 19.6792 |  | 73.3244 | 12.5785 |  | 73.3761 | 254002 |  | 73.4716 | 518.159 |  | 73.5107 | 4727.99 |  | 73.6107 | 4.6828e+19 |  | 73.7073 | 5.4578e+19 |  | 73.7798 | 0
| 73.8948 | 19.4763 |  | 73.989 | 18.6648 |  | 74.0408 | 250427 |  | 74.1362 | 491.582 |  | 74.1753 | 4596.18 |  | 74.2754 | 5.1255e+19 |  | 74.372 | 4.4728e+19 |  | 74.4444 | 0
| 74.5594 | 18.0562 |  | 74.6537 | 16.2303 |  | 74.7055 | 246738 |  | 74.8009 | 474.133 |  | 74.84 | 4684.72 |  | 74.9401 | 4.25e+19 |  | 75.0367 | 4.9254e+19 |  | 75.1091 | 0
| 75.2241 | 15.8245 |  | 75.3184 | 14.4044 |  | 75.3701 | 251218 |  | 75.4656 | 517.956 |  | 75.5047 | 4664.06 |  | 75.6047 | 4.5451e+19 |  | 75.7013 | 4.2332e+19 |  | 75.7738 | 0
| 75.8888 | 15.013 |  | 75.9831 | 13.39 |  | 76.0348 | 248206 |  | 76.1303 | 488.132 |  | 76.1693 | 4510.64 |  | 76.2694 | 4.2106e+19 |  | 76.366 | 3.7539e+19 |  | 76.4384 | 0
| 76.5534 | 15.8245 |  | 76.6477 | 15.013 |  | 76.6995 | 241328 |  | 76.7949 | 483.263 |  | 76.834 | 4365.09 |  | 76.9341 | 4.1909e+19 |  | 77.0307 | 3.7539e+19 |  | 77.1031 | 0
| 77.2181 | 17.8533 |  | 77.3124 | 13.1871 |  | 77.3642 | 243844 |  | 77.4596 | 486.103 |  | 77.4987 | 4415.25 |  | 77.5987 | 4.004e+19 |  | 77.6953 | 3.9669e+19 |  | 77.7678 | 0
| 77.8828 | 13.9986 |  | 77.9771 | 10.7526 |  | 78.0288 | 246574 |  | 78.1243 | 534.593 |  | 78.1634 | 4713.23 |  | 78.2634 | 3.8663e+19 |  | 78.36 | 3.6741e+19 |  | 78.4325 | 0
| 78.5475 | 17.8533 |  | 78.6418 | 16.0274 |  | 78.6935 | 244330 |  | 78.7889 | 491.987 |  | 78.828 | 4647.34 |  | 78.9281 | 3.4137e+19 |  | 79.0247 | 3.2481e+19 |  | 79.0971 | 0
| 79.2121 | 18.6648 |  | 79.3064 | 13.1871 |  | 79.3582 | 241646 |  | 79.4536 | 480.829 |  | 79.4927 | 4533.26 |  | 79.5928 | 3.0891e+19 |  | 79.6894 | 3.1948e+19 |  | 79.7618 | 0
| 79.8768 | 15.2159 |  | 79.9711 | 12.3756 |  | 80.0228 | 247595 |  | 80.1183 | 479.409 |  | 80.1574 | 4455.57 |  | 80.2574 | 3.1088e+19 |  | 80.354 | 3.2481e+19 |  | 80.4265 | 0
| 80.5415 | 19.8821 |  | 80.6358 | 13.9986 |  | 80.6875 | 242428 |  | 80.783 | 474.539 |  | 80.8221 | 4546.04 |  | 80.9221 | 3.0792e+19 |  | 81.0187 | 3.2481e+19 |  | 81.0911 | 0
| 81.2061 | 18.0562 |  | 81.3004 | 17.2447 |  | 81.3522 | 249682 |  | 81.4476 | 485.9 |  | 81.4867 | 4511.63 |  | 81.5868 | 2.9022e+19 |  | 81.6834 | 3.2481e+19 |  | 81.7558 | 0
| 81.8708 | 13.9986 |  | 81.9651 | 11.1583 |  | 82.0169 | 240905 |  | 82.1123 | 488.741 |  | 82.1514 | 4368.05 |  | 82.2514 | 2.971e+19 |  | 82.348 | 2.902e+19 |  | 82.4205 | 0
| 82.5355 | 18.8677 |  | 82.6298 | 14.4044 |  | 82.6815 | 240683 |  | 82.777 | 463.583 |  | 82.8161 | 4521.46 |  | 82.9161 | 2.6562e+19 |  | 83.0127 | 2.6091e+19 |  | 83.0852 | 0
| 83.2002 | 15.4188 |  | 83.2945 | 9.7382 |  | 83.3462 | 235287 |  | 83.4417 | 482.249 |  | 83.4808 | 4441.81 |  | 83.5808 | 2.2922e+19 |  | 83.6774 | 2.4494e+19 |  | 83.7498 | 0
| 83.8648 | 14.8101 |  | 83.9591 | 12.3756 |  | 84.0109 | 232785 |  | 84.1063 | 492.799 |  | 84.1454 | 4269.71 |  | 84.2455 | 2.0955e+19 |  | 84.3421 | 2.2364e+19 |  | 84.4145 | 0
| 84.5295 | 19.0706 |  | 84.6238 | 14.6073 |  | 84.6756 | 234626 |  | 84.771 | 455.468 |  | 84.8101 | 4360.17 |  | 84.9101 | 2.0561e+19 |  | 85.0067 | 2.3162e+19 |  | 85.0792 | 0
| 85.1942 | 17.8533 |  | 85.2885 | 14.8101 |  | 85.3402 | 237104 |  | 85.4357 | 466.829 |  | 85.4748 | 4228.38 |  | 85.5748 | 2.0168e+19 |  | 85.6714 | 2.476e+19 |  | 85.7439 | 0
| 85.8588 | 16.8389 |  | 85.9531 | 11.1583 |  | 86.0049 | 234278 |  | 86.1003 | 454.454 |  | 86.1394 | 4285.44 |  | 86.2395 | 2.2135e+19 |  | 86.3361 | 1.8104e+19 |  | 86.4085 | 0
| 86.5235 | 16.4332 |  | 86.6178 | 15.8245 |  | 86.6696 | 233451 |  | 86.765 | 466.221 |  | 86.8041 | 4278.56 |  | 86.9042 | 2.0069e+19 |  | 87.0007 | 1.6773e+19 |  | 87.0732 | 0
| 87.1882 | 16.6361 |  | 87.2825 | 14.2015 |  | 87.3342 | 228378 |  | 87.4297 | 447.961 |  | 87.4688 | 4312.96 |  | 87.5688 | 1.6823e+19 |  | 87.6654 | 1.6773e+19 |  | 87.7379 | 0
| 87.8529 | 16.4332 |  | 87.9472 | 12.1727 |  | 87.9989 | 230146 |  | 88.0944 | 431.731 |  | 88.1335 | 4179.21 |  | 88.2335 | 1.4954e+19 |  | 88.3301 | 1.4909e+19 |  | 88.4025 | 0
| 88.5175 | 14.8101 |  | 88.6118 | 13.39 |  | 88.6636 | 231004 |  | 88.759 | 481.031 |  | 88.7981 | 4177.24 |  | 88.8982 | 1.7118e+19 |  | 88.9948 | 1.4377e+19 |  | 89.0672 | 0
| 89.1822 | 18.8677 |  | 89.2765 | 15.8245 |  | 89.3282 | 226348 |  | 89.4237 | 441.063 |  | 89.4628 | 4155.59 |  | 89.5628 | 1.4265e+19 |  | 89.6594 | 1.8903e+19 |  | 89.7319 | 0
| 89.8469 | 16.4332 |  | 89.9412 | 11.1583 |  | 89.9929 | 228112 |  | 90.0884 | 461.352 |  | 90.1275 | 4131 |  | 90.2275 | 1.3675e+19 |  | 90.3241 | 1.4377e+19 |  | 90.3966 | 0
| 90.5116 | 16.636 |  | 90.6058 | 15.4188 |  | 90.6576 | 223353 |  | 90.753 | 464.395 |  | 90.7921 | 4064.14 |  | 90.8922 | 1.3773e+19 |  | 90.9888 | 1.3046e+19 |  | 91.0612 | 0
| 91.1762 | 16.4332 |  | 91.2705 | 14.6073 |  | 91.3223 | 229090 |  | 91.4177 | 437.817 |  | 91.4568 | 4072.99 |  | 91.5569 | 1.151e+19 |  | 91.6535 | 1.3312e+19 |  | 91.7259 | 0
| 91.8409 | 18.0562 |  | 91.9352 | 17.2447 |  | 91.9869 | 224533 |  | 92.0824 | 444.715 |  | 92.1215 | 4190.03 |  | 92.2215 | 1.1314e+19 |  | 92.3181 | 9.5845e+18 |  | 92.3906 | 0
| 92.5056 | 14.8101 |  | 92.5999 | 12.9842 |  | 92.6516 | 217420 |  | 92.7471 | 441.266 |  | 92.7862 | 3907.78 |  | 92.8862 | 1.0526e+19 |  | 92.9828 | 1.411e+19 |  | 93.0552 | 0.33635
| 93.1702 | 11.3612 |  | 93.2645 | 16.2303 |  | 93.3163 | 214493 |  | 93.4117 | 415.906 |  | 93.4508 | 4015.96 |  | 93.5509 | 1.0428e+19 |  | 93.6475 | 1.2247e+19 |  | 93.7199 | 0
| 93.8349 | 13.5929 |  | 93.9292 | 11.1583 |  | 93.981 | 213937 |  | 94.0764 | 416.514 |  | 94.1155 | 3920.57 |  | 94.2155 | 9.7395e+18 |  | 94.3121 | 9.5845e+18 |  | 94.3846 | 0
| 94.4996 | 11.7669 |  | 94.5939 | 13.39 |  | 94.6456 | 211928 |  | 94.7411 | 414.08 |  | 94.7802 | 3973.67 |  | 94.8802 | 9.4444e+18 |  | 94.9768 | 1.2779e+19 |  | 95.0493 | 0
| 95.1643 | 16.0274 |  | 95.2586 | 11.3612 |  | 95.3103 | 212476 |  | 95.4058 | 407.385 |  | 95.4448 | 4082.83 |  | 95.5449 | 7.7719e+18 |  | 95.6415 | 8.7858e+18 |  | 95.7139 | 0
| 95.8289 | 17.6504 |  | 95.9232 | 13.5929 |  | 95.975 | 211426 |  | 96.0704 | 405.153 |  | 96.1095 | 3996.28 |  | 96.2096 | 6.5914e+18 |  | 96.3062 | 1.0117e+19 |  | 96.3786 | 0
| 96.4936 | 15.013 |  | 96.5879 | 12.3756 |  | 96.6396 | 210879 |  | 96.7351 | 400.69 |  | 96.7742 | 3890.08 |  | 96.8742 | 8.4606e+18 |  | 96.9708 | 6.3897e+18 |  | 97.0433 | 0
| 97.1583 | 15.013 |  | 97.2526 | 12.5785 |  | 97.3043 | 209401 |  | 97.3998 | 406.776 |  | 97.4389 | 3835.01 |  | 97.5389 | 6.9849e+18 |  | 97.6355 | 5.8572e+18 |  | 97.7079 | 0
| 97.8229 | 14.2015 |  | 97.9172 | 13.9986 |  | 97.969 | 202280 |  | 98.0644 | 400.284 |  | 98.1035 | 3802.55 |  | 98.2036 | 6.8865e+18 |  | 98.3002 | 8.5195e+18 |  | 98.3726 | 0
| 98.4876 | 13.1871 |  | 98.5819 | 10.9554 |  | 98.6337 | 200050 |  | 98.7291 | 384.256 |  | 98.7682 | 3780.92 |  | 98.8683 | 5.9027e+18 |  | 98.9648 | 4.7922e+18 |  | 99.0373 | 0
| 99.1523 | 16.2303 |  | 99.2466 | 10.7526 |  | 99.2983 | 198199 |  | 99.3938 | 386.082 |  | 99.4329 | 3751.42 |  | 99.5329 | 5.3124e+18 |  | 99.6295 | 6.3897e+18 |  | 99.702 | 0
| 99.817 | 15.6216 |  | 99.9113 | 13.7957 |  | 99.963 | 196286 |  | 100.058 | 386.691 |  | 100.098 | 3798.62 |  | 100.198 | 4.427e+18 |  | 100.294 | 5.8572e+18 |  | 100.367 | 0
| 100.482 | 12.7813 |  | 100.576 | 13.5929 |  | 100.628 | 192557 |  | 100.723 | 367.823 |  | 100.762 | 3727.82 |  | 100.862 | 4.2303e+18 |  | 100.959 | 5.0585e+18 |  | 101.031 | 0
| 101.146 | 14.8101 |  | 101.241 | 8.9267 |  | 101.292 | 185627 |  | 101.388 | 354.229 |  | 101.427 | 3702.25 |  | 101.527 | 5.0173e+18 |  | 101.624 | 5.8572e+18 |  | 101.696 | 0
| 101.811 | 12.1727 |  | 101.905 | 11.3612 |  | 101.957 | 184962 |  | 102.052 | 351.795 |  | 102.092 | 3636.35 |  | 102.192 | 4.1319e+18 |  | 102.288 | 4.2598e+18 |  | 102.361 | 0
| 102.476 | 13.9986 |  | 102.57 | 10.7526 |  | 102.622 | 185641 |  | 102.717 | 353.215 |  | 102.756 | 3490.81 |  | 102.856 | 3.9351e+18 |  | 102.953 | 5.3247e+18 |  | 103.025 | 0
| 103.14 | 15.013 |  | 103.235 | 11.5641 |  | 103.286 | 178726 |  | 103.382 | 345.708 |  | 103.421 | 3465.23 |  | 103.521 | 2.853e+18 | [6 empty] | 103.69 | 0
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 14_AsinSiO2_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -44615
Y(um) |  |  | 15455
Z(um) |  |  | 1822
Window coordinates
X(um) |  |  | 1384
Y(um) |  |  | 154
Sample coordinates
X(um) |  |  | 1384
Y(um) |  |  | 154
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-14
Creation time |  | 16:14
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 14_AsinSiO2_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 103.7
Depth calib method | Depth
Mean SR (nm/min) | 0.8625
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method | DOSE
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 500000000000000
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si
Conc calib method | DOSE
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 500000000000000
[C] (at/cm3)
Ref signal | 28Si
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal

## Data_09_As-2
*** DATA FILES ***
Raw data file name(s) | 09_As-2_NEG GOOD.ckb
Sample ID |  | No2 calib sample
Sample code
Data process ID
Analysis date |  | 2017-09-14
Analysis time |  | 10:24
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
09_As-2_NEG GOOD.ckb
As-2 12C Int. (Counts/s) |  |  | As-2 13C Int. (Counts/s) |  |  | As-2 16O Int. (Counts/s) |  |  | As-2 18O Int. (Counts/s) |  |  | As-2 28Si Int. (Counts/s) |  |  | As-2 75As Int. (Counts/s) |  |  | As-2 75As 28Si Int. (Counts/s) |  |  | As-2 197Au Int. (Counts/s) |  |  | As-3 As Conc. using RSF (At./cm3) |  |  | As-2 As Conc. from 103AsSi RSF (At./cm3) |  |  | As-2 As Conc. using SF (At./cm3) |  |  | As-2 As Conc. from 103AsSi SF (At./cm3)
Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth |  | Time[s] | Depth |  | Time[s] | Depth |  | Time[s] | Depth
8.08 |  | 2.2317 | 14.64 |  | 0.20288 | 18.24 |  | 0.51126 | 24.88 |  | 0 | 27.6 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 46.32 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0
54.32 |  | 2.2317 | 60.88 |  | 0.40576 | 64.48 |  | 21.9843 | 71.12 |  | 0.20288 | 73.84 |  | 0 | 80.8 |  | 0 | 87.52 |  | 0 | 92.56 |  | 0 | 80.8 |  | 0 | 87.52 |  | 0 | 80.8 |  | 0 | 87.52 |  | 0
100.56 |  | 3.8547 | 107.12 |  | 0.20288 | 110.72 |  | 166.673 | 117.36 |  | 0.40576 | 120.08 |  | 3.9329 | 127.04 |  | 0.40576 | 133.76 |  | 0.20288 | 138.8 |  | 0.33635 | 127.04 |  | 6.4378e+17 | 133.76 |  | 8.7113e+17 | 127.04 |  | 1.9673e+17 | 133.76 |  | 2.6622e+17
146.8 |  | 2.4345 | 153.36 |  | 1.0144 | 156.96 |  | 244.387 | 163.6 |  | 1.2173 | 166.32 |  | 0 | 173.28 |  | 1.2173 | 180 |  | 0 | 185.04 |  | 0 | 173.28 |  | 1.9314e+18 | 180 |  | 0 | 173.28 |  | 5.9018e+17 | 180 |  | 0
193.04 |  | 3.8547 | 199.6 |  | 0.20288 | 203.2 |  | 452.478 | 209.84 |  | 0.60864 | 212.56 |  | 0.98323 | 219.52 |  | 0.40576 | 226.24 |  | 0.40576 | 231.28 |  | 0 | 219.52 |  | 6.4378e+17 | 226.24 |  | 1.7423e+18 | 219.52 |  | 1.9673e+17 | 226.24 |  | 5.3245e+17
239.28 |  | 3.6518 | 245.84 |  | 1.8259 | 249.44 |  | 858.449 | 256.08 |  | 1.4202 | 258.8 |  | 3.9329 | 265.76 |  | 0.40576 | 272.48 |  | 0.40576 | 277.52 |  | 0.67271 | 265.76 |  | 6.4378e+17 | 272.48 |  | 1.7423e+18 | 265.76 |  | 1.9673e+17 | 272.48 |  | 5.3245e+17
285.52 |  | 2.8403 | 292.08 |  | 1.0144 | 295.68 |  | 1460.27 | 302.32 |  | 2.8403 | 305.04 |  | 8.849 | 312 |  | 1.8259 | 318.72 |  | 0.20288 | 323.76 |  | 0.33635 | 312 |  | 2.897e+18 | 318.72 |  | 8.7113e+17 | 312 |  | 8.8528e+17 | 318.72 |  | 2.6622e+17
331.76 |  | 6.695 | 338.32 |  | 2.8403 | 341.92 |  | 2221.67 | 348.56 |  | 4.4633 | 351.28 |  | 18.6813 | 358.24 |  | 2.2317 | 364.96 |  | 0.81151 | 370 |  | 0.67271 | 358.24 |  | 3.5408e+18 | 364.96 |  | 3.4845e+18 | 358.24 |  | 1.082e+18 | 364.96 |  | 1.0649e+18
378 |  | 12.3756 | 384.56 |  | 4.6662 | 388.16 |  | 3839.18 | 394.8 |  | 9.3324 | 397.52 |  | 34.4129 | 404.48 |  | 7.1007 | 411.2 |  | 0.40576 | 416.24 |  | 1.3454 | 404.48 |  | 1.1266e+19 | 411.2 |  | 1.7423e+18 | 404.48 |  | 3.4427e+18 | 411.2 |  | 5.3245e+17
424.24 |  | 20.4907 | 430.8 |  | 10.7526 | 434.4 |  | 5813.54 | 441.04 |  | 9.7382 | 443.76 |  | 57.027 | 450.72 |  | 9.5353 | 457.44 |  | 1.623 | 462.48 |  | 1.3454 | 450.72 |  | 1.5129e+19 | 457.44 |  | 6.969e+18 | 450.72 |  | 4.6231e+18 | 457.44 |  | 2.1298e+18
470.48 |  | 26.1713 | 477.04 |  | 14.8101 | 480.64 |  | 8170.22 | 487.28 |  | 14.2015 | 490 |  | 108.155 | 496.96 |  | 23.9397 | 503.68 |  | 4.0576 | 508.72 |  | 1.3454 | 496.96 |  | 3.7983e+19 | 503.68 |  | 1.7423e+19 | 496.96 |  | 1.1607e+19 | 503.68 |  | 5.3245e+18
516.72 |  | 46.865 | 523.28 |  | 21.0994 | 526.88 |  | 13572 | 533.52 |  | 21.708 | 536.24 |  | 292.021 | 543.2 |  | 38.5469 | 549.92 |  | 8.9267 | 554.96 |  | 3.0272 | 543.2 |  | 6.116e+19 | 549.92 |  | 3.833e+19 | 543.2 |  | 1.8689e+19 | 549.92 |  | 1.1714e+19
562.96 |  | 53.3571 | 569.52 |  | 29.8232 | 573.12 |  | 17168.9 | 579.76 |  | 32.6635 | 582.48 |  | 447.377 | 589.44 |  | 57.8205 | 596.16 |  | 7.7094 | 601.2 |  | 9.0816 | 589.44 |  | 9.1739e+19 | 596.16 |  | 3.3103e+19 | 589.44 |  | 2.8034e+19 | 596.16 |  | 1.0117e+19
609.2 |  | 56.4003 | 615.76 |  | 31.0404 | 619.36 |  | 22389.9 | 626 |  | 43.2131 | 628.72 |  | 616.495 | 635.68 |  | 74.2537 | 642.4 |  | 18.4619 | 647.44 |  | 6.7271 | 635.68 |  | 1.1781e+20 | 642.4 |  | 7.9273e+19 | 635.68 |  | 3.6001e+19 | 642.4 |  | 2.4226e+19
655.44 |  | 53.3571 | 662 |  | 28.8088 | 665.6 |  | 29363.2 | 672.24 |  | 60.2551 | 674.96 |  | 814.132 | 681.92 |  | 93.3244 | 688.64 |  | 23.7368 | 693.68 |  | 8.7452 | 681.92 |  | 1.4807e+20 | 688.64 |  | 1.0192e+20 | 681.92 |  | 4.5248e+19 | 688.64 |  | 3.1148e+19
701.68 |  | 52.1398 | 708.24 |  | 30.2289 | 711.84 |  | 35031.8 | 718.48 |  | 57.4147 | 721.2 |  | 1002.92 | 728.16 |  | 103.063 | 734.88 |  | 29.6203 | 739.92 |  | 7.7361 | 728.16 |  | 1.6352e+20 | 734.88 |  | 1.2718e+20 | 728.16 |  | 4.9969e+19 | 734.88 |  | 3.8869e+19
747.92 |  | 44.6333 | 754.48 |  | 31.8519 | 758.08 |  | 39478 | 764.72 |  | 81.7602 | 767.44 |  | 1203.52 | 774.4 |  | 109.555 | 781.12 |  | 34.8951 | 786.16 |  | 11.0997 | 774.4 |  | 1.7382e+20 | 781.12 |  | 1.4984e+20 | 774.4 |  | 5.3117e+19 | 781.12 |  | 4.5791e+19
794.16 |  | 52.3427 | 800.72 |  | 30.8375 | 804.32 |  | 54524 | 810.96 |  | 98.3963 | 813.68 |  | 1612.57 | 820.64 |  | 119.699 | 827.36 |  | 36.1124 | 832.4 |  | 7.7362 | 820.64 |  | 1.8992e+20 | 827.36 |  | 1.5506e+20 | 820.64 |  | 5.8035e+19 | 827.36 |  | 4.7388e+19
840.4 |  | 47.2707 | 846.96 |  | 27.3886 | 850.56 |  | 63240.6 | 857.2 |  | 114.221 | 859.92 |  | 1900.69 | 866.88 |  | 128.828 | 873.6 |  | 35.3009 | 878.64 |  | 8.7452 | 866.88 |  | 2.044e+20 | 873.6 |  | 1.5158e+20 | 866.88 |  | 6.2462e+19 | 873.6 |  | 4.6323e+19
886.64 |  | 31.6491 | 893.2 |  | 17.4475 | 896.8 |  | 67632.6 | 903.44 |  | 113.004 | 906.16 |  | 1859.4 | 913.12 |  | 100.019 | 919.84 |  | 38.7498 | 924.88 |  | 3.3635 | 913.12 |  | 1.5869e+20 | 919.84 |  | 1.6639e+20 | 913.12 |  | 4.8494e+19 | 919.84 |  | 5.0849e+19
932.88 |  | 18.6648 | 939.44 |  | 8.9267 | 943.04 |  | 69537.4 | 949.68 |  | 123.351 | 952.4 |  | 1823.99 | 959.36 |  | 75.6738 | 966.08 |  | 35.3009 | 971.12 |  | 2.3545 | 959.36 |  | 1.2007e+20 | 966.08 |  | 1.5158e+20 | 959.36 |  | 3.669e+19 | 966.08 |  | 4.6323e+19
979.12 |  | 10.5497 | 985.68 |  | 5.6806 | 989.28 |  | 67966.6 | 995.92 |  | 123.756 | 998.64 |  | 1886.93 | 1005.6 |  | 41.3872 | 1012.32 |  | 15.2159 | 1017.36 |  | 1.0091 | 1005.6 |  | 6.5666e+19 | 1012.32 |  | 6.5335e+19 | 1005.6 |  | 2.0066e+19 | 1012.32 |  | 1.9967e+19
1025.36 |  | 5.6806 | 1031.92 |  | 5.4777 | 1035.52 |  | 66194.1 | 1042.16 |  | 129.843 | 1044.88 |  | 1725.66 | 1051.84 |  | 23.1282 | 1058.56 |  | 10.9554 | 1063.6 |  | 1.0091 | 1051.84 |  | 3.6696e+19 | 1058.56 |  | 4.7041e+19 | 1051.84 |  | 1.1214e+19 | 1058.56 |  | 1.4376e+19
1071.6 |  | 8.1151 | 1078.16 |  | 5.072 | 1081.76 |  | 63652.7 | 1088.4 |  | 124.568 | 1091.12 |  | 1699.11 | 1098.08 |  | 8.318 | 1104.8 |  | 2.8403 | 1109.84 |  | 0 | 1098.08 |  | 1.3198e+19 | 1104.8 |  | 1.2196e+19 | 1098.08 |  | 4.0329e+18 | 1104.8 |  | 3.7272e+18
1117.84 |  | 4.6662 | 1124.4 |  | 3.4489 | 1128 |  | 60833.3 | 1134.64 |  | 109.758 | 1137.36 |  | 1551.6 | 1144.32 |  | 3.246 | 1151.04 |  | 0.81151 | 1156.08 |  | 0 | 1144.32 |  | 5.1503e+18 | 1151.04 |  | 3.4845e+18 | 1144.32 |  | 1.5738e+18 | 1151.04 |  | 1.0649e+18
1164.08 |  | 4.4633 | 1170.64 |  | 3.6518 | 1174.24 |  | 55338.7 | 1180.88 |  | 108.337 | 1183.6 |  | 1477.86 | 1190.56 |  | 1.8259 | 1197.28 |  | 0.20288 | 1202.32 |  | 0 | 1190.56 |  | 2.897e+18 | 1197.28 |  | 8.7113e+17 | 1190.56 |  | 8.8528e+17 | 1197.28 |  | 2.6622e+17
1210.32 |  | 2.2317 | 1216.88 |  | 4.4633 | 1220.48 |  | 52211.4 | 1227.12 |  | 98.3964 | 1229.84 |  | 1360.84 | 1236.8 |  | 0 | 1243.52 |  | 0 | 1248.56 |  | 0 | 1236.8 |  | 0 | 1243.52 |  | 0 | 1236.8 |  | 0 | 1243.52 |  | 0
1256.56 |  | 2.8403 | 1263.12 |  | 1.8259 | 1266.72 |  | 51694.7 | 1273.36 |  | 92.7157 | 1276.08 |  | 1429.67 | 1283.04 |  | 0.20288 | 1289.76 |  | 0 | 1294.8 |  | 0 | 1283.04 |  | 3.2189e+17 | 1289.76 |  | 0 | 1283.04 |  | 98364100000000000 | 1289.76 |  | 0
1302.8 |  | 4.2604 | 1309.36 |  | 3.4489 | 1312.96 |  | 49926.7 | 1319.6 |  | 96.9762 | 1322.32 |  | 1376.58 | 1329.28 |  | 0 | 1336 |  | 0 | 1341.04 |  | 0 | 1329.28 |  | 0 | 1336 |  | 0 | 1329.28 |  | 0 | 1336 |  | 0
1349.04 |  | 5.2748 | 1355.6 |  | 1.0144 | 1359.2 |  | 47453.3 | 1365.84 |  | 83.789 | 1368.56 |  | 1290.04 | 1375.52 |  | 0 | 1382.24 |  | 0 | 1387.28 |  | 0 | 1375.52 |  | 0 | 1382.24 |  | 0 | 1375.52 |  | 0 | 1382.24 |  | 0
1395.28 |  | 4.2604 | 1401.84 |  | 3.6518 | 1405.44 |  | 45323.4 | 1412.08 |  | 91.4984 | 1414.8 |  | 1328.4 | 1421.76 |  | 0 | 1428.48 |  | 0 | 1433.52 |  | 0 | 1421.76 |  | 0 | 1428.48 |  | 0 | 1421.76 |  | 0 | 1428.48 |  | 0
1441.52 |  | 4.8691 | 1448.08 |  | 2.8403 | 1451.68 |  | 45364.8 | 1458.32 |  | 90.8898 | 1461.04 |  | 1278.25 | 1468 |  | 0 | 1474.72 |  | 0 | 1479.76 |  | 0 | 1468 |  | 0 | 1474.72 |  | 0 | 1468 |  | 0 | 1474.72 |  | 0
1487.76 |  | 3.8547 | 1494.32 |  | 3.4489 | 1497.92 |  | 46414.6 | 1504.56 |  | 93.1215 | 1507.28 |  | 1277.26 | 1514.24 |  | 0 | 1520.96 |  | 0 | 1526 |  | 0 | 1514.24 |  | 0 | 1520.96 |  | 0 | 1514.24 |  | 0 | 1520.96 |  | 0
1534 |  | 3.6518 | 1540.56 |  | 3.8547 | 1544.16 |  | 46309.5 | 1550.8 |  | 94.9475 | 1553.52 |  | 1244.81 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1572.24 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0
1580.24 |  | 3.4489 | 1586.8 |  | 3.0432 | 1590.4 |  | 45540.4 | 1597.04 |  | 89.8754 | 1599.76 |  | 1249.73 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1618.48 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0
1626.48 |  | 4.8691 | 1633.04 |  | 4.0576 | 1636.64 |  | 50578.5 | 1643.28 |  | 87.6437 | 1646 |  | 1245.8 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1664.72 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0
1672.72 |  | 2.8403 | 1679.28 |  | 2.0288 | 1682.88 |  | 51984 | 1689.52 |  | 96.7733 | 1692.24 |  | 1359.86 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1710.96 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0
1718.96 |  | 3.246 | 1725.52 |  | 3.246 | 1729.12 |  | 45428.5 | 1735.76 |  | 83.3833 | 1738.48 |  | 1266.45 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1757.2 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0
1765.2 |  | 4.2604 | 1771.76 |  | 2.0288 | 1775.36 |  | 48301.4 | 1782 |  | 83.7891 | 1784.72 |  | 1260.55 | 1791.68 |  | 0.20288 | 1798.4 |  | 0 | 1803.44 |  | 0 | 1791.68 |  | 3.2189e+17 | 1798.4 |  | 0 | 1791.68 |  | 98364100000000000 | 1798.4 |  | 0
1811.44 |  | 5.072 | 1818 |  | 3.6518 | 1821.6 |  | 48559.4 | 1828.24 |  | 89.2668 | 1830.96 |  | 1255.63 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1849.68 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0
1857.68 |  | 4.6662 | 1864.24 |  | 1.8259 | 1867.84 |  | 52794.3 | 1874.48 |  | 97.5848 | 1877.2 |  | 1308.73 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1895.92 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0
1903.92 |  | 4.0576 | 1910.48 |  | 2.6374 | 1914.08 |  | 47736 | 1920.72 |  | 86.8322 | 1923.44 |  | 1270.38 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1942.16 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0
1950.16 |  | 4.8691 | 1956.72 |  | 3.0432 | 1960.32 |  | 50966.4 | 1966.96 |  | 99.005 | 1969.68 |  | 1396.25 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1988.4 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0
1996.4 |  | 4.4633 | 2002.96 |  | 4.2604 | 2006.56 |  | 50007.2 | 2013.2 |  | 92.1071 | 2015.92 |  | 1376.58 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2034.64 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0
2042.64 |  | 2.8403 | 2049.2 |  | 3.4489 | 2052.8 |  | 49821.1 | 2059.44 |  | 95.1503 | 2062.16 |  | 1331.35 | 2069.12 |  | 0 | 2075.84 |  | 0 | 2080.88 |  | 0 | 2069.12 |  | 0 | 2075.84 |  | 0 | 2069.12 |  | 0 | 2075.84 |  | 0
2088.88 |  | 3.4489 | 2095.44 |  | 3.246 | 2099.04 |  | 48233.7 | 2105.68 |  | 95.9618 | 2108.4 |  | 1297.92 | 2115.36 |  | 0.20288 | 2122.08 |  | 0.20288 | 2127.12 |  | 0 | 2115.36 |  | 3.2189e+17 | 2122.08 |  | 8.7113e+17 | 2115.36 |  | 98364100000000000 | 2122.08 |  | 2.6622e+17
2135.12 |  | 4.2604 | 2141.68 |  | 2.6374 | 2145.28 |  | 48152.3 | 2151.92 |  | 93.3244 | 2154.64 |  | 1293.98 | 2161.6 |  | 0 | 2168.32 |  | 0 | 2173.36 |  | 0 | 2161.6 |  | 0 | 2168.32 |  | 0 | 2161.6 |  | 0 | 2168.32 |  | 0
2181.36 |  | 5.2748 | 2187.92 |  | 3.0432 | 2191.52 |  | 49758.4 | 2198.16 |  | 94.5417 | 2200.88 |  | 1272.34 | 2207.84 |  | 0.20288 | 2214.56 |  | 0 | 2219.6 |  | 0 | 2207.84 |  | 3.2189e+17 | 2214.56 |  | 0 | 2207.84 |  | 98364100000000000 | 2214.56 |  | 0
2227.6 |  | 2.8403 | 2234.16 |  | 4.4633 | 2237.76 |  | 52710.1 | 2244.4 |  | 86.8322 | 2247.12 |  | 1365.76 | 2254.08 |  | 0.20288 | 2260.8 |  | 0 | 2265.84 |  | 0 | 2254.08 |  | 3.2189e+17 | 2260.8 |  | 0 | 2254.08 |  | 98364100000000000 | 2260.8 |  | 0
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 09_As-2_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -30180
Y(um) |  |  | -13965
Z(um) |  |  | 1896
Window coordinates
X(um) |  |  | 1320
Y(um) |  |  | -265
Sample coordinates
X(um) |  |  | 1320
Y(um) |  |  | -265
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-14
Creation time |  | 10:24
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 09_As-2_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 1000
Depth calib method | Depth
Mean SR (nm/min) | 26.48
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal

## Data_07_As-1
*** DATA FILES ***
Raw data file name(s) | 07_As-1_NEG GOOD.ckb
Sample ID |  | GaAs calib sample
Sample code
Data process ID
Analysis date |  | 2017-09-13
Analysis time |  | 16:58
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
07_As-1_NEG GOOD.ckb
As-1 12C Int. (Counts/s) |  |  | As-1 13C Int. (Counts/s) |  |  | As-1 16O Int. (Counts/s) |  |  | As-1 18O Int. (Counts/s) |  |  | As-1 28Si Int. (Counts/s) |  |  | As-1 75As Int. (Counts/s) |  |  | As-1 75As 28Si Int. (Counts/s) |  |  | As-1 197Au Int. (Counts/s) |  |  | As-3 As Conc. using RSF (At./cm3) |  |  | As-2 As Conc. from 103AsSi RSF (At./cm3) |  |  | As-1 As Conc. using SF (At./cm3) |  |  | As-1 As Conc. from 103AsSi SF (At./cm3)
Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s]
8.08 |  | 2.8403 | 14.64 |  | 0 | 18.24 |  | 2.045 | 24.88 |  | 0 | 27.6 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 46.32 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0
54.32 |  | 1.8259 | 60.88 |  | 0 | 64.48 |  | 16.3604 | 71.12 |  | 0 | 73.84 |  | 0 | 80.8 |  | 0.60864 | 87.52 |  | 0.20288 | 92.56 |  | 0.33635 | 80.8 |  | 6.8409e+17 | 87.52 |  | 6.1717e+17 | 80.8 |  | 2.9509e+17 | 87.52 |  | 2.6622e+17
100.56 |  | 3.8547 | 107.12 |  | 0.81151 | 110.72 |  | 450.945 | 117.36 |  | 2.2317 | 120.08 |  | 2.9497 | 127.04 |  | 0.81151 | 133.76 |  | 0 | 138.8 |  | 0 | 127.04 |  | 9.1212e+17 | 133.76 |  | 0 | 127.04 |  | 3.9346e+17 | 133.76 |  | 0
146.8 |  | 3.0432 | 153.36 |  | 0.81151 | 156.96 |  | 892.702 | 163.6 |  | 1.2173 | 166.32 |  | 3.9329 | 173.28 |  | 1.4202 | 180 |  | 0.20288 | 185.04 |  | 0 | 173.28 |  | 1.5962e+18 | 180 |  | 6.1717e+17 | 173.28 |  | 6.8855e+17 | 180 |  | 2.6622e+17
193.04 |  | 6.2892 | 199.6 |  | 1.623 | 203.2 |  | 1174.94 | 209.84 |  | 2.8403 | 212.56 |  | 0.98323 | 219.52 |  | 2.0288 | 226.24 |  | 0.40576 | 231.28 |  | 0.33635 | 219.52 |  | 2.2803e+18 | 226.24 |  | 1.2343e+18 | 219.52 |  | 9.8364e+17 | 226.24 |  | 5.3245e+17
239.28 |  | 9.5353 | 245.84 |  | 2.0288 | 249.44 |  | 1929.65 | 256.08 |  | 2.8403 | 258.8 |  | 8.849 | 265.76 |  | 1.623 | 272.48 |  | 0.20288 | 277.52 |  | 0.67271 | 265.76 |  | 1.8242e+18 | 272.48 |  | 6.1717e+17 | 265.76 |  | 7.8691e+17 | 272.48 |  | 2.6622e+17
285.52 |  | 9.7382 | 292.08 |  | 3.8547 | 295.68 |  | 3076.64 | 302.32 |  | 7.9123 | 305.04 |  | 26.5471 | 312 |  | 2.6374 | 318.72 |  | 0.40576 | 323.76 |  | 0.67271 | 312 |  | 2.9644e+18 | 318.72 |  | 1.2343e+18 | 312 |  | 1.2787e+18 | 318.72 |  | 5.3245e+17
331.76 |  | 13.1871 | 338.32 |  | 3.6518 | 341.92 |  | 4301.94 | 348.56 |  | 8.5209 | 351.28 |  | 31.4632 | 358.24 |  | 5.072 | 364.96 |  | 1.0144 | 370 |  | 1.3454 | 358.24 |  | 5.7008e+18 | 364.96 |  | 3.0858e+18 | 358.24 |  | 2.4591e+18 | 364.96 |  | 1.3311e+18
378 |  | 22.9253 | 384.56 |  | 9.5353 | 388.16 |  | 6760.01 | 394.8 |  | 10.7526 | 397.52 |  | 63.9096 | 404.48 |  | 10.3468 | 411.2 |  | 1.0144 | 416.24 |  | 0 | 404.48 |  | 1.163e+19 | 411.2 |  | 3.0858e+18 | 404.48 |  | 5.0166e+18 | 411.2 |  | 1.3311e+18
424.24 |  | 33.2721 | 430.8 |  | 13.5929 | 434.4 |  | 10767.1 | 441.04 |  | 21.5051 | 443.76 |  | 124.87 | 450.72 |  | 20.8965 | 457.44 |  | 3.0432 | 462.48 |  | 1.6818 | 450.72 |  | 2.3487e+19 | 457.44 |  | 9.2575e+18 | 450.72 |  | 1.0132e+19 | 457.44 |  | 3.9934e+18
470.48 |  | 58.2262 | 477.04 |  | 23.9397 | 480.64 |  | 17810.5 | 487.28 |  | 31.2433 | 490 |  | 280.221 | 496.96 |  | 36.5182 | 503.68 |  | 4.2604 | 508.72 |  | 3.0272 | 496.96 |  | 4.1045e+19 | 503.68 |  | 1.2961e+19 | 496.96 |  | 1.7706e+19 | 503.68 |  | 5.5907e+18
516.72 |  | 86.2235 | 523.28 |  | 46.2563 | 526.88 |  | 24931.6 | 533.52 |  | 48.488 | 536.24 |  | 463.105 | 543.2 |  | 58.0233 | 549.92 |  | 7.9123 | 554.96 |  | 5.718 | 543.2 |  | 6.5217e+19 | 549.92 |  | 2.407e+19 | 543.2 |  | 2.8132e+19 | 549.92 |  | 1.0383e+19
562.96 |  | 88.6581 | 569.52 |  | 48.6909 | 573.12 |  | 34146.8 | 579.76 |  | 72.4278 | 582.48 |  | 684.338 | 589.44 |  | 92.3099 | 596.16 |  | 14.6072 | 601.2 |  | 6.0544 | 589.44 |  | 1.0375e+20 | 596.16 |  | 4.4436e+19 | 589.44 |  | 4.4756e+19 | 596.16 |  | 1.9168e+19
609.2 |  | 95.3531 | 615.76 |  | 47.0679 | 619.36 |  | 42678.8 | 626 |  | 82.3689 | 628.72 |  | 1039.3 | 635.68 |  | 108.946 | 642.4 |  | 19.6792 | 647.44 |  | 13.4541 | 635.68 |  | 1.2245e+20 | 642.4 |  | 5.9865e+19 | 635.68 |  | 5.2822e+19 | 642.4 |  | 2.5824e+19
655.44 |  | 84.8033 | 662 |  | 52.5456 | 665.6 |  | 54747.8 | 672.24 |  | 93.3243 | 674.96 |  | 1238.91 | 681.92 |  | 133.292 | 688.64 |  | 27.7944 | 693.68 |  | 9.4179 | 681.92 |  | 1.4982e+20 | 688.64 |  | 8.4552e+19 | 681.92 |  | 6.4625e+19 | 688.64 |  | 3.6473e+19
701.68 |  | 79.9342 | 708.24 |  | 54.7772 | 711.84 |  | 64968.9 | 718.48 |  | 126.799 | 721.2 |  | 1680.42 | 728.16 |  | 152.362 | 734.88 |  | 31.6491 | 739.92 |  | 11.0997 | 728.16 |  | 1.7125e+20 | 734.88 |  | 9.6278e+19 | 728.16 |  | 7.3872e+19 | 734.88 |  | 4.1531e+19
747.92 |  | 77.2969 | 754.48 |  | 47.2708 | 758.08 |  | 74214.5 | 764.72 |  | 140.19 | 767.44 |  | 1887.91 | 774.4 |  | 150.131 | 781.12 |  | 40.17 | 786.16 |  | 9.0816 | 774.4 |  | 1.6874e+20 | 781.12 |  | 1.222e+20 | 774.4 |  | 7.279e+19 | 781.12 |  | 5.2713e+19
794.16 |  | 79.7314 | 800.72 |  | 38.1412 | 804.32 |  | 88524 | 810.96 |  | 149.116 | 813.68 |  | 2269.44 | 820.64 |  | 149.928 | 827.36 |  | 40.7786 | 832.4 |  | 8.0725 | 820.64 |  | 1.6852e+20 | 827.36 |  | 1.2405e+20 | 820.64 |  | 7.2691e+19 | 827.36 |  | 5.3511e+19
840.4 |  | 50.5168 | 846.96 |  | 32.8663 | 850.56 |  | 97400.4 | 857.2 |  | 188.069 | 859.92 |  | 2481.85 | 866.88 |  | 141.001 | 873.6 |  | 45.8506 | 878.64 |  | 6.3907 | 866.88 |  | 1.5848e+20 | 873.6 |  | 1.3948e+20 | 866.88 |  | 6.8363e+19 | 873.6 |  | 6.0167e+19
886.64 |  | 37.5325 | 893.2 |  | 22.3166 | 896.8 |  | 104633 | 903.44 |  | 203.894 | 906.16 |  | 2562.47 | 913.12 |  | 111.989 | 919.84 |  | 30.8375 | 924.88 |  | 3.0272 | 913.12 |  | 1.2587e+20 | 919.84 |  | 9.381e+19 | 913.12 |  | 5.4297e+19 | 919.84 |  | 4.0466e+19
932.88 |  | 25.5627 | 939.44 |  | 18.2591 | 943.04 |  | 105672 | 949.68 |  | 183.809 | 952.4 |  | 2544.78 | 959.36 |  | 71.6162 | 966.08 |  | 23.1282 | 971.12 |  | 1.3454 | 959.36 |  | 8.0495e+19 | 966.08 |  | 7.0357e+19 | 959.36 |  | 3.4723e+19 | 966.08 |  | 3.035e+19
979.12 |  | 20.085 | 985.68 |  | 13.1871 | 989.28 |  | 105874 | 995.92 |  | 193.141 | 998.64 |  | 2510.35 | 1005.6 |  | 41.1844 | 1012.32 |  | 11.5641 | 1017.36 |  | 2.0181 | 1005.6 |  | 4.629e+19 | 1012.32 |  | 3.5179e+19 | 1005.6 |  | 1.9968e+19 | 1012.32 |  | 1.5175e+19
1025.36 |  | 13.39 | 1031.92 |  | 14.4044 | 1035.52 |  | 100988 | 1042.16 |  | 185.635 | 1044.88 |  | 2420.88 | 1051.84 |  | 18.8677 | 1058.56 |  | 4.8691 | 1063.6 |  | 0.67271 | 1051.84 |  | 2.1207e+19 | 1058.56 |  | 1.4812e+19 | 1051.84 |  | 9.1479e+18 | 1058.56 |  | 6.3894e+18
1071.6 |  | 13.5929 | 1078.16 |  | 13.39 | 1081.76 |  | 101797 | 1088.4 |  | 184.012 | 1091.12 |  | 2393.34 | 1098.08 |  | 11.1583 | 1104.8 |  | 3.6518 | 1109.84 |  | 0.33635 | 1098.08 |  | 1.2542e+19 | 1104.8 |  | 1.1109e+19 | 1098.08 |  | 5.41e+18 | 1104.8 |  | 4.792e+18
1117.84 |  | 13.39 | 1124.4 |  | 11.5641 | 1128 |  | 98380.1 | 1134.64 |  | 174.882 | 1137.36 |  | 2271.4 | 1144.32 |  | 5.4777 | 1151.04 |  | 1.2173 | 1156.08 |  | 0 | 1144.32 |  | 6.1568e+18 | 1151.04 |  | 3.703e+18 | 1144.32 |  | 2.6558e+18 | 1151.04 |  | 1.5974e+18
1164.08 |  | 11.9698 | 1170.64 |  | 8.5209 | 1174.24 |  | 98935.5 | 1180.88 |  | 171.027 | 1183.6 |  | 2250.76 | 1190.56 |  | 1.623 | 1197.28 |  | 0.20288 | 1202.32 |  | 0 | 1190.56 |  | 1.8242e+18 | 1197.28 |  | 6.1717e+17 | 1190.56 |  | 7.8691e+17 | 1197.28 |  | 2.6622e+17
1210.32 |  | 8.7238 | 1216.88 |  | 8.7238 | 1220.48 |  | 96368 | 1227.12 |  | 177.925 | 1229.84 |  | 2133.74 | 1236.8 |  | 1.623 | 1243.52 |  | 0.40576 | 1248.56 |  | 0 | 1236.8 |  | 1.8242e+18 | 1243.52 |  | 1.2343e+18 | 1236.8 |  | 7.8691e+17 | 1243.52 |  | 5.3245e+17
1256.56 |  | 13.39 | 1263.12 |  | 9.7382 | 1266.72 |  | 95139.5 | 1273.36 |  | 165.55 | 1276.08 |  | 2268.46 | 1283.04 |  | 0.40576 | 1289.76 |  | 0.20288 | 1294.8 |  | 0 | 1283.04 |  | 4.5606e+17 | 1289.76 |  | 6.1717e+17 | 1283.04 |  | 1.9673e+17 | 1289.76 |  | 2.6622e+17
1302.8 |  | 8.7238 | 1309.36 |  | 9.1295 | 1312.96 |  | 95677 | 1319.6 |  | 174.071 | 1322.32 |  | 2147.51 | 1329.28 |  | 0.20288 | 1336 |  | 0 | 1341.04 |  | 0 | 1329.28 |  | 2.2803e+17 | 1336 |  | 0 | 1329.28 |  | 98364100000000000 | 1336 |  | 0
1349.04 |  | 11.3612 | 1355.6 |  | 8.318 | 1359.2 |  | 96744.4 | 1365.84 |  | 160.275 | 1368.56 |  | 2072.78 | 1375.52 |  | 0.40576 | 1382.24 |  | 0 | 1387.28 |  | 0 | 1375.52 |  | 4.5606e+17 | 1382.24 |  | 0 | 1375.52 |  | 1.9673e+17 | 1382.24 |  | 0
1395.28 |  | 9.7382 | 1401.84 |  | 9.3324 | 1405.44 |  | 97262.8 | 1412.08 |  | 167.984 | 1414.8 |  | 2124.89 | 1421.76 |  | 0.20288 | 1428.48 |  | 0 | 1433.52 |  | 0 | 1421.76 |  | 2.2803e+17 | 1428.48 |  | 0 | 1421.76 |  | 98364100000000000 | 1428.48 |  | 0
1441.52 |  | 8.318 | 1448.08 |  | 10.3468 | 1451.68 |  | 90634.8 | 1458.32 |  | 170.825 | 1461.04 |  | 2075.72 | 1468 |  | 0.20288 | 1474.72 |  | 0 | 1479.76 |  | 0 | 1468 |  | 2.2803e+17 | 1474.72 |  | 0 | 1468 |  | 98364100000000000 | 1474.72 |  | 0
1487.76 |  | 9.5353 | 1494.32 |  | 9.5353 | 1497.92 |  | 91362.4 | 1504.56 |  | 163.724 | 1507.28 |  | 2029.5 | 1514.24 |  | 0.20288 | 1520.96 |  | 0 | 1526 |  | 0 | 1514.24 |  | 2.2803e+17 | 1520.96 |  | 0 | 1514.24 |  | 98364100000000000 | 1520.96 |  | 0
1534 |  | 8.318 | 1540.56 |  | 8.9267 | 1544.16 |  | 88698.8 | 1550.8 |  | 159.26 | 1553.52 |  | 1931.18 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1572.24 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0
1580.24 |  | 7.5065 | 1586.8 |  | 9.3324 | 1590.4 |  | 87486.8 | 1597.04 |  | 167.376 | 1599.76 |  | 1852.51 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1618.48 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0
1626.48 |  | 10.7526 | 1633.04 |  | 9.9411 | 1636.64 |  | 81583.3 | 1643.28 |  | 140.19 | 1646 |  | 1766.96 | 1652.96 |  | 0.20288 | 1659.68 |  | 0 | 1664.72 |  | 0 | 1652.96 |  | 2.2803e+17 | 1659.68 |  | 0 | 1652.96 |  | 98364100000000000 | 1659.68 |  | 0
1672.72 |  | 9.941 | 1679.28 |  | 9.5353 | 1682.88 |  | 80233.5 | 1689.52 |  | 145.059 | 1692.24 |  | 1762.04 | 1699.2 |  | 0.60864 | 1705.92 |  | 0 | 1710.96 |  | 0 | 1699.2 |  | 6.8409e+17 | 1705.92 |  | 0 | 1699.2 |  | 2.9509e+17 | 1705.92 |  | 0
1718.96 |  | 7.9123 | 1725.52 |  | 8.5209 | 1729.12 |  | 73022.9 | 1735.76 |  | 125.177 | 1738.48 |  | 1647 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1757.2 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0
1765.2 |  | 8.7238 | 1771.76 |  | 8.1151 | 1775.36 |  | 74504.4 | 1782 |  | 127.408 | 1784.72 |  | 1560.46 | 1791.68 |  | 0 | 1798.4 |  | 0 | 1803.44 |  | 0 | 1791.68 |  | 0 | 1798.4 |  | 0 | 1791.68 |  | 0 | 1798.4 |  | 0
1811.44 |  | 7.9123 | 1818 |  | 7.5065 | 1821.6 |  | 72929.4 | 1828.24 |  | 132.277 | 1830.96 |  | 1496.55 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1849.68 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0
1857.68 |  | 6.695 | 1864.24 |  | 5.2748 | 1867.84 |  | 67814 | 1874.48 |  | 123.756 | 1877.2 |  | 1506.38 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1895.92 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0
1903.92 |  | 7.7094 | 1910.48 |  | 5.6806 | 1914.08 |  | 67046.7 | 1920.72 |  | 120.51 | 1923.44 |  | 1492.61 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1942.16 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0
1950.16 |  | 8.318 | 1956.72 |  | 6.2892 | 1960.32 |  | 63043.7 | 1966.96 |  | 109.555 | 1969.68 |  | 1357.91 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1988.4 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0
1996.4 |  | 7.1007 | 2002.96 |  | 5.072 | 2006.56 |  | 65168.6 | 2013.2 |  | 108.337 | 2015.92 |  | 1337.25 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2034.64 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0
2042.64 |  | 7.1007 | 2049.2 |  | 5.072 | 2052.8 |  | 60891.4 | 2059.44 |  | 116.047 | 2062.16 |  | 1352.98 | 2069.12 |  | 0.20288 | 2075.84 |  | 0 | 2080.88 |  | 0 | 2069.12 |  | 2.2803e+17 | 2075.84 |  | 0 | 2069.12 |  | 98364100000000000 | 2075.84 |  | 0
2088.88 |  | 9.9411 | 2095.44 |  | 7.1007 | 2099.04 |  | 58689.3 | 2105.68 |  | 96.5705 | 2108.4 |  | 1328.4 | 2115.36 |  | 0 | 2122.08 |  | 0 | 2127.12 |  | 0 | 2115.36 |  | 0 | 2122.08 |  | 0 | 2115.36 |  | 0 | 2122.08 |  | 0
2135.12 |  | 4.2604 | 2141.68 |  | 5.072 | 2145.28 |  | 57453.5 | 2151.92 |  | 107.932 | 2154.64 |  | 1219.25 | 2161.6 |  | 0.20288 | 2168.32 |  | 0 | 2173.36 |  | 0 | 2161.6 |  | 2.2803e+17 | 2168.32 |  | 0 | 2161.6 |  | 98364100000000000 | 2168.32 |  | 0
2181.36 |  | 6.8979 | 2187.92 |  | 4.8691 | 2191.52 |  | 53691.8 | 2198.16 |  | 91.0927 | 2200.88 |  | 1197.62 | 2207.84 |  | 0 | 2214.56 |  | 0 | 2219.6 |  | 0 | 2207.84 |  | 0 | 2214.56 |  | 0 | 2207.84 |  | 0 | 2214.56 |  | 0
2227.6 |  | 5.2748 | 2234.16 |  | 4.4633 | [5 empty] | 2244.4 |  | 92.3099 | [5 empty] | 2254.08 |  | 0 | 2260.8 |  | 0 | [5 empty] | 2254.08 |  | 0 | 2260.8 |  | 0 | 2254.08 | 0 | 2260.8 | 0
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 07_As-1_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -43250
Y(um) |  |  | -13675
Z(um) |  |  | 1926
Window coordinates
X(um) |  |  | 2750
Y(um) |  |  | 24
Sample coordinates
X(um) |  |  | 2750
Y(um) |  |  | 24
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-13
Creation time |  | 16:58
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 07_As-1_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 1000
Depth calib method | Depth
Mean SR (nm/min) | 26.48
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal

## Data_10_As-3
*** DATA FILES ***
Raw data file name(s) | 10_As-3_NEG GOOD.ckb
Sample ID |  | No 1
Sample code
Data process ID
Analysis date |  | 2017-09-14
Analysis time |  | 12:05
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
10_As-3_NEG GOOD.ckb
As-3 12C Int. (Counts/s) |  |  | As-3 13C Int. (Counts/s) |  |  | As-3 16O Int. (Counts/s) |  |  | As-3 18O Int. (Counts/s) |  |  | As-3 28Si Int. (Counts/s) |  |  | As-3 75As Int. (Counts/s) |  |  | As-3 75As 28Si Int. (Counts/s) |  |  | As-3 197Au Int. (Counts/s) |  |  | As-3 As Conc. using RSF (At./cm3) |  |  | As-2 As Conc. from 103AsSi RSF (At./cm3) |  |  | As-3 As Conc. using SF (At./cm3) |  |  | As-3 As Conc. from 103AsSi SF (At./cm3)
Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s]
8.08 |  | 1.0144 | 14.64 |  | 0 | 18.24 |  | 0 | 24.88 |  | 0 | 27.6 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 46.32 |  | 0.33635 | 34.56 |  | 0 | 41.28 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0
54.32 |  | 2.6374 | 60.88 |  | 0.20288 | 64.48 |  | 14.3154 | 71.12 |  | 0 | 73.84 |  | 0 | 80.8 |  | 0.40576 | 87.52 |  | 0 | 92.56 |  | 0.67271 | 80.8 |  | 8.446e+17 | 87.52 |  | 0 | 80.8 |  | 1.9673e+17 | 87.52 |  | 0
100.56 |  | 2.0288 | 107.12 |  | 0.20288 | 110.72 |  | 96.6293 | 117.36 |  | 0.40576 | 120.08 |  | 1.9665 | 127.04 |  | 0.40576 | 133.76 |  | 0.60864 | 138.8 |  | 0 | 127.04 |  | 8.446e+17 | 133.76 |  | 3.43e+18 | 127.04 |  | 1.9673e+17 | 133.76 |  | 7.9868e+17
146.8 |  | 0.60864 | 153.36 |  | 0.60864 | 156.96 |  | 263.307 | 163.6 |  | 1.2173 | 166.32 |  | 2.9497 | 173.28 |  | 1.0144 | 180 |  | 0.20288 | 185.04 |  | 0 | 173.28 |  | 2.1115e+18 | 180 |  | 1.1433e+18 | 173.28 |  | 4.9182e+17 | 180 |  | 2.6622e+17
193.04 |  | 1.623 | 199.6 |  | 0.20288 | 203.2 |  | 512.301 | 209.84 |  | 1.0144 | 212.56 |  | 2.9497 | 219.52 |  | 1.2173 | 226.24 |  | 0.20288 | 231.28 |  | 0 | 219.52 |  | 2.5338e+18 | 226.24 |  | 1.1433e+18 | 219.52 |  | 5.9018e+17 | 226.24 |  | 2.6622e+17
239.28 |  | 3.0432 | 245.84 |  | 0.60864 | 249.44 |  | 857.946 | 256.08 |  | 3.0432 | 258.8 |  | 4.9161 | 265.76 |  | 2.4345 | 272.48 |  | 0.20288 | 277.52 |  | 0.33635 | 265.76 |  | 5.0676e+18 | 272.48 |  | 1.1433e+18 | 265.76 |  | 1.1804e+18 | 272.48 |  | 2.6622e+17
285.52 |  | 3.6518 | 292.08 |  | 1.0144 | 295.68 |  | 1333.97 | 302.32 |  | 2.2317 | 305.04 |  | 7.8658 | 312 |  | 4.0576 | 318.72 |  | 0.60864 | 323.76 |  | 0.67271 | 312 |  | 8.446e+18 | 318.72 |  | 3.43e+18 | 312 |  | 1.9673e+18 | 318.72 |  | 7.9868e+17
331.76 |  | 5.4777 | 338.32 |  | 2.4345 | 341.92 |  | 1993.64 | 348.56 |  | 2.8403 | 351.28 |  | 14.7484 | 358.24 |  | 5.8835 | 364.96 |  | 0.81151 | 370 |  | 1.0091 | 358.24 |  | 1.2247e+19 | 364.96 |  | 4.5733e+18 | 358.24 |  | 2.8526e+18 | 364.96 |  | 1.0649e+18
378 |  | 9.3324 | 384.56 |  | 3.4489 | 388.16 |  | 3528.27 | 394.8 |  | 8.318 | 397.52 |  | 39.329 | 404.48 |  | 8.1151 | 411.2 |  | 1.4202 | 416.24 |  | 0.33635 | 404.48 |  | 1.6892e+19 | 411.2 |  | 8.0032e+18 | 404.48 |  | 3.9346e+18 | 411.2 |  | 1.8636e+18
424.24 |  | 13.1871 | 430.8 |  | 6.2892 | 434.4 |  | 5630.74 | 441.04 |  | 11.3612 | 443.76 |  | 86.5239 | 450.72 |  | 20.085 | 457.44 |  | 2.6374 | 462.48 |  | 5.718 | 450.72 |  | 4.1808e+19 | 457.44 |  | 1.4863e+19 | 450.72 |  | 9.738e+18 | 457.44 |  | 3.4609e+18
470.48 |  | 25.9685 | 477.04 |  | 9.5353 | 480.64 |  | 8742.08 | 487.28 |  | 20.085 | 490 |  | 147.484 | 496.96 |  | 31.852 | 503.68 |  | 5.4777 | 508.72 |  | 6.3907 | 496.96 |  | 6.6301e+19 | 503.68 |  | 3.087e+19 | 496.96 |  | 1.5443e+19 | 503.68 |  | 7.1881e+18
516.72 |  | 35.7066 | 523.28 |  | 15.4188 | 526.88 |  | 12492.9 | 533.52 |  | 26.78 | 536.24 |  | 295.955 | 543.2 |  | 53.1542 | 549.92 |  | 7.9123 | 554.96 |  | 8.4088 | 543.2 |  | 1.1064e+20 | 549.92 |  | 4.4589e+19 | 543.2 |  | 2.5772e+19 | 549.92 |  | 1.0383e+19
562.96 |  | 40.5757 | 569.52 |  | 23.1282 | 573.12 |  | 16070.6 | 579.76 |  | 34.2865 | 582.48 |  | 498.506 | 589.44 |  | 75.8767 | 596.16 |  | 11.5641 | 601.2 |  | 7.7362 | 589.44 |  | 1.5794e+20 | 596.16 |  | 6.5169e+19 | 589.44 |  | 3.6788e+19 | 596.16 |  | 1.5175e+19
609.2 |  | 34.6922 | 615.76 |  | 18.8677 | 619.36 |  | 20595.6 | 626 |  | 34.6923 | 628.72 |  | 644.031 | 635.68 |  | 86.0208 | 642.4 |  | 15.2159 | 647.44 |  | 11.436 | 635.68 |  | 1.7906e+20 | 642.4 |  | 8.5749e+19 | 635.68 |  | 4.1707e+19 | 642.4 |  | 1.9967e+19
655.44 |  | 33.2721 | 662 |  | 17.4475 | 665.6 |  | 24444.8 | 672.24 |  | 45.4449 | 674.96 |  | 825.943 | 681.92 |  | 106.106 | 688.64 |  | 27.1858 | 693.68 |  | 10.7633 | 681.92 |  | 2.2086e+20 | 688.64 |  | 1.532e+20 | 681.92 |  | 5.1445e+19 | 688.64 |  | 3.5674e+19
701.68 |  | 31.0404 | 708.24 |  | 19.0706 | 711.84 |  | 29566.1 | 718.48 |  | 51.937 | 721.2 |  | 986.216 | 728.16 |  | 107.12 | 734.88 |  | 21.708 | 739.92 |  | 13.1178 | 728.16 |  | 2.2298e+20 | 734.88 |  | 1.2234e+20 | 728.16 |  | 5.1936e+19 | 734.88 |  | 2.8486e+19
747.92 |  | 27.7944 | 754.48 |  | 18.2591 | 758.08 |  | 33920.1 | 764.72 |  | 71.2105 | 767.44 |  | 1064.88 | 774.4 |  | 103.468 | 781.12 |  | 27.3886 | 786.16 |  | 7.7362 | 774.4 |  | 2.1537e+20 | 781.12 |  | 1.5435e+20 | 774.4 |  | 5.0166e+19 | 781.12 |  | 3.594e+19
794.16 |  | 25.5627 | 800.72 |  | 17.8533 | 804.32 |  | 38368.1 | 810.96 |  | 73.2393 | 813.68 |  | 1277.28 | 820.64 |  | 103.671 | 827.36 |  | 25.3598 | 832.4 |  | 12.1087 | 820.64 |  | 2.158e+20 | 827.36 |  | 1.4292e+20 | 820.64 |  | 5.0264e+19 | 827.36 |  | 3.3278e+19
840.4 |  | 19.6792 | 846.96 |  | 12.9842 | 850.56 |  | 42549.6 | 857.2 |  | 82.9776 | 859.92 |  | 1317.59 | 866.88 |  | 99.005 | 873.6 |  | 27.9973 | 878.64 |  | 10.427 | 866.88 |  | 2.0608e+20 | 873.6 |  | 1.5778e+20 | 866.88 |  | 4.8002e+19 | 873.6 |  | 3.6739e+19
886.64 |  | 12.7814 | 893.2 |  | 9.7382 | 896.8 |  | 44255.5 | 903.44 |  | 85.4121 | 906.16 |  | 1297.94 | 913.12 |  | 88.0495 | 919.84 |  | 25.5627 | 924.88 |  | 6.7271 | 913.12 |  | 1.8328e+20 | 919.84 |  | 1.4406e+20 | 913.12 |  | 4.269e+19 | 919.84 |  | 3.3544e+19
932.88 |  | 8.318 | 939.44 |  | 6.4921 | 943.04 |  | 45301.7 | 949.68 |  | 86.6293 | 952.4 |  | 1390.36 | 959.36 |  | 57.6176 | 966.08 |  | 23.331 | 971.12 |  | 3.3635 | 959.36 |  | 1.1993e+20 | 966.08 |  | 1.3148e+20 | 959.36 |  | 2.7936e+19 | 966.08 |  | 3.0616e+19
979.12 |  | 3.8547 | 985.68 |  | 2.6374 | 989.28 |  | 43631.2 | 995.92 |  | 82.5718 | 998.64 |  | 1271.38 | 1005.6 |  | 38.1412 | 1012.32 |  | 13.39 | 1017.36 |  | 0.67271 | 1005.6 |  | 7.9392e+19 | 1012.32 |  | 7.5459e+19 | 1005.6 |  | 1.8492e+19 | 1012.32 |  | 1.7571e+19
1025.36 |  | 4.2604 | 1031.92 |  | 2.4345 | 1035.52 |  | 40963.3 | 1042.16 |  | 76.0796 | 1044.88 |  | 1216.32 | 1051.84 |  | 21.0994 | 1058.56 |  | 5.6806 | 1063.6 |  | 0 | 1051.84 |  | 4.3919e+19 | 1058.56 |  | 3.2013e+19 | 1051.84 |  | 1.023e+19 | 1058.56 |  | 7.4543e+18
1071.6 |  | 1.8259 | 1078.16 |  | 2.8403 | 1081.76 |  | 37323.4 | 1088.4 |  | 76.4854 | 1091.12 |  | 1076.68 | 1098.08 |  | 7.9123 | 1104.8 |  | 4.4633 | 1109.84 |  | 0.33635 | 1098.08 |  | 1.647e+19 | 1104.8 |  | 2.5153e+19 | 1098.08 |  | 3.8362e+18 | 1104.8 |  | 5.857e+18
1117.84 |  | 2.4345 | 1124.4 |  | 2.4345 | 1128 |  | 34975 | 1134.64 |  | 64.9214 | 1137.36 |  | 1075.7 | 1144.32 |  | 3.0432 | 1151.04 |  | 1.8259 | 1156.08 |  | 0.33635 | 1144.32 |  | 6.3345e+18 | 1151.04 |  | 1.029e+19 | 1144.32 |  | 1.4755e+18 | 1151.04 |  | 2.396e+18
1164.08 |  | 1.8259 | 1170.64 |  | 3.0432 | 1174.24 |  | 34969 | 1180.88 |  | 67.153 | 1183.6 |  | 977.371 | 1190.56 |  | 1.2173 | 1197.28 |  | 0.40576 | 1202.32 |  | 0.33635 | 1190.56 |  | 2.5338e+18 | 1197.28 |  | 2.2866e+18 | 1190.56 |  | 5.9018e+17 | 1197.28 |  | 5.3245e+17
1210.32 |  | 1.4202 | 1216.88 |  | 2.0288 | 1220.48 |  | 33197 | 1227.12 |  | 59.8493 | 1229.84 |  | 983.269 | 1236.8 |  | 0.60864 | 1243.52 |  | 0.20288 | 1248.56 |  | 0 | 1236.8 |  | 1.2669e+18 | 1243.52 |  | 1.1433e+18 | 1236.8 |  | 2.9509e+17 | 1243.52 |  | 2.6622e+17
1256.56 |  | 2.2317 | 1263.12 |  | 1.8259 | 1266.72 |  | 32619.7 | 1273.36 |  | 66.1385 | 1276.08 |  | 971.472 | 1283.04 |  | 0.20288 | 1289.76 |  | 0 | 1294.8 |  | 0 | 1283.04 |  | 4.223e+17 | 1289.76 |  | 0 | 1283.04 |  | 98364100000000000 | 1289.76 |  | 0
1302.8 |  | 2.2317 | 1309.36 |  | 1.2173 | 1312.96 |  | 31741.9 | 1319.6 |  | 58.632 | 1322.32 |  | 923.288 | 1329.28 |  | 0 | 1336 |  | 0 | 1341.04 |  | 0 | 1329.28 |  | 0 | 1336 |  | 0 | 1329.28 |  | 0 | 1336 |  | 0
1349.04 |  | 1.0144 | 1355.6 |  | 2.0288 | 1359.2 |  | 34483.3 | 1365.84 |  | 65.1241 | 1368.56 |  | 999.986 | 1375.52 |  | 0.20288 | 1382.24 |  | 0 | 1387.28 |  | 0 | 1375.52 |  | 4.223e+17 | 1382.24 |  | 0 | 1375.52 |  | 98364100000000000 | 1382.24 |  | 0
1395.28 |  | 1.8259 | 1401.84 |  | 1.4202 | 1405.44 |  | 28603.5 | 1412.08 |  | 65.53 | 1414.8 |  | 800.381 | 1421.76 |  | 0 | 1428.48 |  | 0 | 1433.52 |  | 0 | 1421.76 |  | 0 | 1428.48 |  | 0 | 1421.76 |  | 0 | 1428.48 |  | 0
1441.52 |  | 1.0144 | 1448.08 |  | 1.4202 | 1451.68 |  | 30006 | 1458.32 |  | 52.9514 | 1461.04 |  | 890.84 | 1468 |  | 0 | 1474.72 |  | 0 | 1479.76 |  | 0 | 1468 |  | 0 | 1474.72 |  | 0 | 1468 |  | 0 | 1474.72 |  | 0
1487.76 |  | 2.6374 | 1494.32 |  | 2.8403 | 1497.92 |  | 27388.1 | 1504.56 |  | 51.3283 | 1507.28 |  | 861.34 | 1514.24 |  | 0.20288 | 1520.96 |  | 0 | 1526 |  | 0 | 1514.24 |  | 4.223e+17 | 1520.96 |  | 0 | 1514.24 |  | 98364100000000000 | 1520.96 |  | 0
1534 |  | 1.4202 | 1540.56 |  | 1.623 | 1544.16 |  | 26864.1 | 1550.8 |  | 52.9514 | 1553.52 |  | 791.531 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1572.24 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0
1580.24 |  | 2.2317 | 1586.8 |  | 1.2173 | 1590.4 |  | 27367.9 | 1597.04 |  | 44.6334 | 1599.76 |  | 765.957 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1618.48 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0
1626.48 |  | 1.2173 | 1633.04 |  | 1.4202 | 1636.64 |  | 27403.7 | 1643.28 |  | 51.1255 | 1646 |  | 797.43 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1664.72 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0
1672.72 |  | 1.2173 | 1679.28 |  | 1.623 | 1682.88 |  | 27966.3 | 1689.52 |  | 48.2851 | 1692.24 |  | 890.84 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1710.96 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0
1718.96 |  | 1.2173 | 1725.52 |  | 1.2173 | 1729.12 |  | 26613.8 | 1735.76 |  | 48.6909 | 1738.48 |  | 734.496 | 1745.44 |  | 0 | 1752.16 |  | 0.20288 | 1757.2 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 1.1433e+18 | 1745.44 |  | 0 | 1752.16 |  | 2.6622e+17
1765.2 |  | 1.8259 | 1771.76 |  | 1.4202 | 1775.36 |  | 28239.1 | 1782 |  | 49.9082 | 1784.72 |  | 789.561 | 1791.68 |  | 0 | 1798.4 |  | 0 | 1803.44 |  | 0 | 1791.68 |  | 0 | 1798.4 |  | 0 | 1791.68 |  | 0 | 1798.4 |  | 0
1811.44 |  | 1.2173 | 1818 |  | 1.4202 | 1821.6 |  | 28534.5 | 1828.24 |  | 52.7485 | 1830.96 |  | 840.699 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1849.68 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0 | 1837.92 |  | 0 | 1844.64 |  | 0
1857.68 |  | 0.81151 | 1864.24 |  | 1.4202 | 1867.84 |  | 27902.3 | 1874.48 |  | 57.4148 | 1877.2 |  | 760.065 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1895.92 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0 | 1884.16 |  | 0 | 1890.88 |  | 0
1903.92 |  | 1.4202 | 1910.48 |  | 1.8259 | 1914.08 |  | 29687.6 | 1920.72 |  | 63.2983 | 1923.44 |  | 807.265 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1942.16 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0 | 1930.4 |  | 0 | 1937.12 |  | 0
1950.16 |  | 1.623 | 1956.72 |  | 0.60864 | 1960.32 |  | 28861.7 | 1966.96 |  | 60.0522 | 1969.68 |  | 872.164 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1988.4 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0 | 1976.64 |  | 0 | 1983.36 |  | 0
1996.4 |  | 1.2173 | 2002.96 |  | 1.0144 | 2006.56 |  | 26881.2 | 2013.2 |  | 46.6621 | 2015.92 |  | 896.743 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2034.64 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0 | 2022.88 |  | 0 | 2029.6 |  | 0
2042.64 |  | 1.623 | 2049.2 |  | 0.81151 | 2052.8 |  | 28347.2 | 2059.44 |  | 55.7918 | 2062.16 |  | 766.952 | 2069.12 |  | 0 | 2075.84 |  | 0 | 2080.88 |  | 0 | 2069.12 |  | 0 | 2075.84 |  | 0 | 2069.12 |  | 0 | 2075.84 |  | 0
2088.88 |  | 1.8259 | 2095.44 |  | 1.623 | 2099.04 |  | 31757.8 | 2105.68 |  | 63.907 | 2108.4 |  | 902.646 | 2115.36 |  | 0.20288 | 2122.08 |  | 0 | 2127.12 |  | 0 | 2115.36 |  | 4.223e+17 | 2122.08 |  | 0 | 2115.36 |  | 98364100000000000 | 2122.08 |  | 0
2135.12 |  | 2.2317 | 2141.68 |  | 2.2317 | 2145.28 |  | 32803.7 | 2151.92 |  | 64.3128 | 2154.64 |  | 878.064 | 2161.6 |  | 0 | 2168.32 |  | 0 | 2173.36 |  | 0 | 2161.6 |  | 0 | 2168.32 |  | 0 | 2161.6 |  | 0 | 2168.32 |  | 0
2181.36 |  | 0.81151 | 2187.92 |  | 1.8259 | 2191.52 |  | 28809.2 | 2198.16 |  | 58.0234 | 2200.88 |  | 808.244 | 2207.84 |  | 0 | 2214.56 |  | 0 | 2219.6 |  | 0 | 2207.84 |  | 0 | 2214.56 |  | 0 | 2207.84 |  | 0 | 2214.56 |  | 0
2227.6 |  | 2.6374 | 2234.16 |  | 1.8259 | 2237.76 |  | 31572.2 | 2244.4 |  | 62.2839 | 2247.12 |  | 896.747 | 2254.08 |  | 0 | 2260.8 |  | 0.20288 | 2265.84 |  | 0 | 2254.08 |  | 0 | 2260.8 |  | 1.1433e+18 | 2254.08 |  | 0 | 2260.8 |  | 2.6622e+17
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 10_As-3_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -43515
Y(um) |  |  | 1715
Z(um) |  |  | 1886
Window coordinates
X(um) |  |  | 2484
Y(um) |  |  | 915
Sample coordinates
X(um) |  |  | 2484
Y(um) |  |  | 915
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-14
Creation time |  | 12:05
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 10_As-3_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 1000
Depth calib method | Depth
Mean SR (nm/min) | 26.48
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal

## Data_11_As-1
*** DATA FILES ***
Raw data file name(s) | 11_As-1_NEG GOOD.ckb
Sample ID |  | GaAs calib sample
Sample code
Data process ID
Analysis date |  | 2017-09-14
Analysis time |  | 13:53
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
11_As-1_NEG GOOD.ckb
As-1 Bis 12C Int. (Counts/s) |  |  | As-1 Bis 13C Int. (Counts/s) |  |  | As-1 Bis 16O Int. (Counts/s) |  |  | As-1 Bis 18O Int. (Counts/s) |  |  | As-1 Bis 28Si Int. (Counts/s) |  |  | As-1 Bis 75As Int. (Counts/s) |  |  | As-1 Bis 75As 28Si Int. (Counts/s) |  |  | As-1 Bis 197Au Int. (Counts/s) |  |  | As-3 As Conc. using RSF (At./cm3) |  |  | As-2 As Conc. from 103AsSi RSF (At./cm3) |  |  | As-1 Bis As Conc. using SF (At./cm3) |  |  | As-1 Bis As Conc. from 103AsSi SF (At./cm3)
Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s]
8.08 |  | 2.6374 | 14.64 |  | 0 | 18.24 |  | 1.0225 | 24.88 |  | 0 | 27.6 |  | 0 | 34.56 |  | 0.20288 | 41.28 |  | 0.20288 | 46.32 |  | 0 | 34.56 |  | 8.6015e+17 | 41.28 |  | 2.3264e+18 | 34.56 |  | 98364100000000000 | 41.28 |  | 2.6622e+17
54.32 |  | 2.0288 | 60.88 |  | 0 | 64.48 |  | 9.2027 | 71.12 |  | 0 | 73.84 |  | 0 | 80.8 |  | 1.0144 | 87.52 |  | 0 | 92.56 |  | 0 | 80.8 |  | 4.3008e+18 | 87.52 |  | 0 | 80.8 |  | 4.9182e+17 | 87.52 |  | 0
100.56 |  | 2.0288 | 107.12 |  | 0 | 110.72 |  | 76.6899 | 117.36 |  | 0 | 120.08 |  | 0 | 127.04 |  | 1.2173 | 133.76 |  | 0.40576 | 138.8 |  | 0 | 127.04 |  | 5.1609e+18 | 133.76 |  | 4.6528e+18 | 127.04 |  | 5.9018e+17 | 133.76 |  | 5.3245e+17
146.8 |  | 1.623 | 153.36 |  | 0.40576 | 156.96 |  | 138.554 | 163.6 |  | 0.60864 | 166.32 |  | 2.9497 | 173.28 |  | 1.4202 | 180 |  | 0.40576 | 185.04 |  | 0.33635 | 173.28 |  | 6.0211e+18 | 180 |  | 4.6528e+18 | 173.28 |  | 6.8855e+17 | 180 |  | 5.3245e+17
193.04 |  | 2.6374 | 199.6 |  | 0.81151 | 203.2 |  | 224.959 | 209.84 |  | 0.81151 | 212.56 |  | 2.9497 | 219.52 |  | 2.0288 | 226.24 |  | 0.40576 | 231.28 |  | 0.33635 | 219.52 |  | 8.6015e+18 | 226.24 |  | 4.6528e+18 | 219.52 |  | 9.8364e+17 | 226.24 |  | 5.3245e+17
239.28 |  | 1.2173 | 245.84 |  | 0 | 249.44 |  | 338.463 | 256.08 |  | 1.4202 | 258.8 |  | 5.8994 | 265.76 |  | 1.2173 | 272.48 |  | 0.81151 | 277.52 |  | 0 | 265.76 |  | 5.1609e+18 | 272.48 |  | 9.3055e+18 | 265.76 |  | 5.9018e+17 | 272.48 |  | 1.0649e+18
285.52 |  | 1.2173 | 292.08 |  | 0.81151 | 295.68 |  | 616.093 | 302.32 |  | 1.0144 | 305.04 |  | 9.8323 | 312 |  | 2.2317 | 318.72 |  | 0.81151 | 323.76 |  | 0.33635 | 312 |  | 9.4617e+18 | 318.72 |  | 9.3055e+18 | 312 |  | 1.082e+18 | 318.72 |  | 1.0649e+18
331.76 |  | 1.8259 | 338.32 |  | 0.20288 | 341.92 |  | 845.163 | 348.56 |  | 2.2317 | 351.28 |  | 11.7987 | 358.24 |  | 5.6806 | 364.96 |  | 0.60864 | 370 |  | 1.3454 | 358.24 |  | 2.4084e+19 | 364.96 |  | 6.9792e+18 | 358.24 |  | 2.7542e+18 | 364.96 |  | 7.9868e+17
378 |  | 3.6518 | 384.56 |  | 1.4202 | 388.16 |  | 1292.56 | 394.8 |  | 2.6374 | 397.52 |  | 21.6309 | 404.48 |  | 7.3036 | 411.2 |  | 1.4202 | 416.24 |  | 0.33635 | 404.48 |  | 3.0966e+19 | 411.2 |  | 1.6285e+19 | 404.48 |  | 3.5411e+18 | 411.2 |  | 1.8636e+18
424.24 |  | 4.4633 | 430.8 |  | 2.2317 | 434.4 |  | 2179.22 | 441.04 |  | 2.2317 | 443.76 |  | 45.2286 | 450.72 |  | 11.5641 | 457.44 |  | 2.6374 | 462.48 |  | 3.0272 | 450.72 |  | 4.9029e+19 | 457.44 |  | 3.0243e+19 | 450.72 |  | 5.6068e+18 | 457.44 |  | 3.4609e+18
470.48 |  | 6.4921 | 477.04 |  | 2.4345 | 480.64 |  | 2621.09 | 487.28 |  | 7.3036 | 490 |  | 52.1109 | 496.96 |  | 21.9109 | 503.68 |  | 3.0432 | 508.72 |  | 5.3817 | 496.96 |  | 9.2897e+19 | 503.68 |  | 3.4896e+19 | 496.96 |  | 1.0623e+19 | 503.68 |  | 3.9934e+18
516.72 |  | 9.3324 | 523.28 |  | 3.6518 | 526.88 |  | 3585.07 | 533.52 |  | 7.9123 | 536.24 |  | 107.172 | 543.2 |  | 30.4318 | 549.92 |  | 8.7238 | 554.96 |  | 5.718 | 543.2 |  | 1.2902e+20 | 549.92 |  | 1.0004e+20 | 543.2 |  | 1.4755e+19 | 549.92 |  | 1.1448e+19
562.96 |  | 10.3468 | 569.52 |  | 4.4633 | 573.12 |  | 5115.36 | 579.76 |  | 17.2447 | 582.48 |  | 198.613 | 589.44 |  | 53.7629 | 596.16 |  | 8.318 | 601.2 |  | 9.7543 | 589.44 |  | 2.2794e+20 | 596.16 |  | 9.5382e+19 | 589.44 |  | 2.6066e+19 | 596.16 |  | 1.0915e+19
609.2 |  | 13.39 | 615.76 |  | 5.4777 | 619.36 |  | 6312.25 | 626 |  | 14.8101 | 628.72 |  | 293.988 | 635.68 |  | 65.5299 | 642.4 |  | 15.2159 | 647.44 |  | 8.4088 | 635.68 |  | 2.7783e+20 | 642.4 |  | 1.7448e+20 | 635.68 |  | 3.1772e+19 | 642.4 |  | 1.9967e+19
655.44 |  | 10.5497 | 662 |  | 4.4633 | 665.6 |  | 7491.92 | 672.24 |  | 16.8389 | 674.96 |  | 348.065 | 681.92 |  | 76.8911 | 688.64 |  | 20.2879 | 693.68 |  | 7.7361 | 681.92 |  | 3.26e+20 | 688.64 |  | 2.3264e+20 | 681.92 |  | 3.728e+19 | 688.64 |  | 2.6622e+19
701.68 |  | 8.318 | 708.24 |  | 6.8979 | 711.84 |  | 7973.53 | 718.48 |  | 16.636 | 721.2 |  | 377.563 | 728.16 |  | 82.3688 | 734.88 |  | 24.1426 | 739.92 |  | 11.436 | 728.16 |  | 3.4922e+20 | 734.88 |  | 2.7684e+20 | 728.16 |  | 3.9936e+19 | 734.88 |  | 3.1681e+19
747.92 |  | 8.318 | 754.48 |  | 6.0864 | 758.08 |  | 9731.26 | 764.72 |  | 21.708 | 767.44 |  | 483.757 | 774.4 |  | 87.8466 | 781.12 |  | 26.5771 | 786.16 |  | 14.4632 | 774.4 |  | 3.7245e+20 | 781.12 |  | 3.0476e+20 | 774.4 |  | 4.2592e+19 | 781.12 |  | 3.4876e+19
794.16 |  | 7.9123 | 800.72 |  | 3.8547 | 804.32 |  | 11484.5 | 810.96 |  | 27.1857 | 813.68 |  | 606.662 | 820.64 |  | 101.642 | 827.36 |  | 34.2865 | 832.4 |  | 10.427 | 820.64 |  | 4.3094e+20 | 827.36 |  | 3.9316e+20 | 820.64 |  | 4.9281e+19 | 827.36 |  | 4.4992e+19
840.4 |  | 4.4633 | 846.96 |  | 6.0864 | 850.56 |  | 13417 | 857.2 |  | 28.2002 | 859.92 |  | 640.096 | 866.88 |  | 97.7877 | 873.6 |  | 32.2577 | 878.64 |  | 10.7633 | 866.88 |  | 4.146e+20 | 873.6 |  | 3.699e+20 | 866.88 |  | 4.7412e+19 | 873.6 |  | 4.233e+19
886.64 |  | 3.6518 | 893.2 |  | 4.0576 | 896.8 |  | 14326.7 | 903.44 |  | 32.0548 | 906.16 |  | 663.693 | 913.12 |  | 88.861 | 919.84 |  | 30.2289 | 924.88 |  | 9.4179 | 913.12 |  | 3.7675e+20 | 919.84 |  | 3.4663e+20 | 913.12 |  | 4.3084e+19 | 919.84 |  | 3.9668e+19
932.88 |  | 4.2604 | 939.44 |  | 2.4345 | 943.04 |  | 15057.6 | 949.68 |  | 31.2433 | 952.4 |  | 717.776 | 959.36 |  | 69.7903 | 966.08 |  | 23.7368 | 971.12 |  | 4.7089 | 959.36 |  | 2.9589e+20 | 966.08 |  | 2.7219e+20 | 959.36 |  | 3.3837e+19 | 966.08 |  | 3.1148e+19
979.12 |  | 2.0288 | 985.68 |  | 1.2173 | 989.28 |  | 14431.9 | 995.92 |  | 29.2145 | 998.64 |  | 735.472 | 1005.6 |  | 43.416 | 1012.32 |  | 20.2879 | 1017.36 |  | 4.7089 | 1005.6 |  | 1.8407e+20 | 1012.32 |  | 2.3264e+20 | 1005.6 |  | 2.105e+19 | 1012.32 |  | 2.6622e+19
1025.36 |  | 0.81151 | 1031.92 |  | 1.0144 | 1035.52 |  | 13370.2 | 1042.16 |  | 25.1569 | 1044.88 |  | 670.579 | 1051.84 |  | 26.1713 | 1058.56 |  | 11.3612 | 1063.6 |  | 2.0181 | 1051.84 |  | 1.1096e+20 | 1058.56 |  | 1.3028e+20 | 1051.84 |  | 1.2689e+19 | 1058.56 |  | 1.4909e+19
1071.6 |  | 0.81151 | 1078.16 |  | 0.40576 | 1081.76 |  | 11610.7 | 1088.4 |  | 27.5915 | 1091.12 |  | 570.283 | 1098.08 |  | 17.8533 | 1104.8 |  | 8.5209 | 1109.84 |  | 1.0091 | 1098.08 |  | 7.5694e+19 | 1104.8 |  | 9.7708e+19 | 1098.08 |  | 8.656e+18 | 1104.8 |  | 1.1182e+19
1117.84 |  | 0.81151 | 1124.4 |  | 0.40576 | 1128 |  | 11087.7 | 1134.64 |  | 21.5051 | 1137.36 |  | 521.119 | 1144.32 |  | 8.318 | 1151.04 |  | 3.0432 | 1156.08 |  | 0.33635 | 1144.32 |  | 3.5266e+19 | 1151.04 |  | 3.4896e+19 | 1144.32 |  | 4.0329e+18 | 1151.04 |  | 3.9934e+18
1164.08 |  | 0.81151 | 1170.64 |  | 0.60864 | 1174.24 |  | 9540.43 | 1180.88 |  | 25.5627 | 1183.6 |  | 529.973 | 1190.56 |  | 3.8547 | 1197.28 |  | 1.4202 | 1202.32 |  | 0.33635 | 1190.56 |  | 1.6343e+19 | 1197.28 |  | 1.6285e+19 | 1190.56 |  | 1.8689e+18 | 1197.28 |  | 1.8636e+18
1210.32 |  | 0.60864 | 1216.88 |  | 0 | 1220.48 |  | 9185.23 | 1227.12 |  | 21.3022 | 1229.84 |  | 434.592 | 1236.8 |  | 1.8259 | 1243.52 |  | 0.60864 | 1248.56 |  | 0 | 1236.8 |  | 7.7414e+18 | 1243.52 |  | 6.9792e+18 | 1236.8 |  | 8.8528e+17 | 1243.52 |  | 7.9868e+17
1256.56 |  | 0.20288 | 1263.12 |  | 0.20288 | 1266.72 |  | 8242.23 | 1273.36 |  | 19.4763 | 1276.08 |  | 467.042 | 1283.04 |  | 1.2173 | 1289.76 |  | 1.2173 | 1294.8 |  | 0 | 1283.04 |  | 5.1609e+18 | 1289.76 |  | 1.3958e+19 | 1283.04 |  | 5.9018e+17 | 1289.76 |  | 1.5974e+18
1302.8 |  | 0.40576 | 1309.36 |  | 0 | 1312.96 |  | 8267.46 | 1319.6 |  | 20.4907 | 1322.32 |  | 445.409 | 1329.28 |  | 0.60864 | 1336 |  | 0 | 1341.04 |  | 0 | 1329.28 |  | 2.5805e+18 | 1336 |  | 0 | 1329.28 |  | 2.9509e+17 | 1336 |  | 0
1349.04 |  | 0.20288 | 1355.6 |  | 0.81151 | 1359.2 |  | 7817.6 | 1365.84 |  | 20.085 | 1368.56 |  | 389.362 | 1375.52 |  | 0.81151 | 1382.24 |  | 0.20288 | 1387.28 |  | 0 | 1375.52 |  | 3.4406e+18 | 1382.24 |  | 2.3264e+18 | 1375.52 |  | 3.9346e+17 | 1382.24 |  | 2.6622e+17
1395.28 |  | 0.40576 | 1401.84 |  | 0.81151 | 1405.44 |  | 7266.24 | 1412.08 |  | 17.2447 | 1414.8 |  | 410.012 | 1421.76 |  | 0 | 1428.48 |  | 0.20288 | 1433.52 |  | 0 | 1421.76 |  | 0 | 1428.48 |  | 2.3264e+18 | 1421.76 |  | 0 | 1428.48 |  | 2.6622e+17
1441.52 |  | 0.60864 | 1448.08 |  | 0.20288 | 1451.68 |  | 7375.59 | 1458.32 |  | 16.8389 | 1461.04 |  | 345.119 | 1468 |  | 0.20288 | 1474.72 |  | 0 | 1479.76 |  | 0 | 1468 |  | 8.6015e+17 | 1474.72 |  | 0 | 1468 |  | 98364100000000000 | 1474.72 |  | 0
1487.76 |  | 0.20288 | 1494.32 |  | 0.40576 | 1497.92 |  | 7164.19 | 1504.56 |  | 16.8389 | 1507.28 |  | 394.28 | 1514.24 |  | 0.20288 | 1520.96 |  | 0.20288 | 1526 |  | 0 | 1514.24 |  | 8.6015e+17 | 1520.96 |  | 2.3264e+18 | 1514.24 |  | 98364100000000000 | 1520.96 |  | 2.6622e+17
1534 |  | 0.20288 | 1540.56 |  | 0.40576 | 1544.16 |  | 6571.83 | 1550.8 |  | 14.6073 | 1553.52 |  | 358.884 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1572.24 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0 | 1560.48 |  | 0 | 1567.2 |  | 0
1580.24 |  | 0.20288 | 1586.8 |  | 0.20288 | 1590.4 |  | 6134.14 | 1597.04 |  | 12.9842 | 1599.76 |  | 346.1 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1618.48 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0
1626.48 |  | 0.20288 | 1633.04 |  | 0 | 1636.64 |  | 6082.54 | 1643.28 |  | 12.5785 | 1646 |  | 300.87 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1664.72 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0 | 1652.96 |  | 0 | 1659.68 |  | 0
1672.72 |  | 0.40576 | 1679.28 |  | 0.40576 | 1682.88 |  | 5728.7 | 1689.52 |  | 10.1439 | 1692.24 |  | 284.157 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1710.96 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0
1718.96 |  | 0.40576 | 1725.52 |  | 0.40576 | 1729.12 |  | 5219.5 | 1735.76 |  | 12.7813 | 1738.48 |  | 312.67 | 1745.44 |  | 0.20288 | 1752.16 |  | 0 | 1757.2 |  | 0 | 1745.44 |  | 8.6015e+17 | 1752.16 |  | 0 | 1745.44 |  | 98364100000000000 | 1752.16 |  | 0
1765.2 |  | 0 | 1771.76 |  | 0 | 1775.36 |  | 5241.35 | 1782 |  | 15.2159 | 1784.72 |  | 283.173 | 1791.68 |  | 0.20288 | 1798.4 |  | 0.20288 | 1803.44 |  | 0 | 1791.68 |  | 8.6015e+17 | 1798.4 |  | 2.3264e+18 | 1791.68 |  | 98364100000000000 | 1798.4 |  | 2.6622e+17
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 11_As-1_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -42705
Y(um) |  |  | -13780
Z(um) |  |  | 1812
Window coordinates
X(um) |  |  | 3295
Y(um) |  |  | -80
Sample coordinates
X(um) |  |  | 3295
Y(um) |  |  | -80
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-14
Creation time |  | 13:53
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 11_As-1_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 1000
Depth calib method | Depth
Mean SR (nm/min) | 33.27
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal

## Data_12_As-2
*** DATA FILES ***
Raw data file name(s) | 12_As-2_NEG GOOD.ckb
Sample ID |  | No2 calib sample
Sample code
Data process ID
Analysis date |  | 2017-09-14
Analysis time |  | 14:34
Analysis recipe name | 5 new mass calib now with eGun.rckb
*** DATA START ***
12_As-2_NEG GOOD.ckb
As-2 Bis 12C Int. (Counts/s) |  |  | As-2 Bis 13C Int. (Counts/s) |  |  | As-2 Bis 16O Int. (Counts/s) |  |  | As-2 Bis 18O Int. (Counts/s) |  |  | As-2 Bis 28Si Int. (Counts/s) |  |  | As-2 Bis 75As Int. (Counts/s) |  |  | As-2 Bis 75As 28Si Int. (Counts/s) |  |  | As-2 Bis 197Au Int. (Counts/s) |  |  | As-3 As Conc. using RSF (At./cm3) |  |  | As-2 As Conc. from 103AsSi RSF (At./cm3) |  |  | As-2 Bis As Conc. using SF (At./cm3) |  |  | As-2 Bis As Conc. from 103AsSi SF (At./cm3)
Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s] | Time[s] | Depth | I[cnt/s]
8.08 |  | 3.6518 | 14.64 |  | 0.20288 | 18.24 |  | 0 | 24.88 |  | 0 | 27.6 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 46.32 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0 | 34.56 |  | 0 | 41.28 |  | 0
54.32 |  | 3.0432 | 60.88 |  | 1.0144 | 64.48 |  | 203.487 | 71.12 |  | 1.0144 | 73.84 |  | 5.8994 | 80.8 |  | 1.623 | 87.52 |  | 0.20288 | 92.56 |  | 0 | 80.8 |  | 1.6541e+18 | 87.52 |  | 5.5959e+17 | 80.8 |  | 7.8691e+17 | 87.52 |  | 2.6622e+17
100.56 |  | 21.0994 | 107.12 |  | 5.8835 | 110.72 |  | 2875.77 | 117.36 |  | 4.8691 | 120.08 |  | 20.6477 | 127.04 |  | 4.2604 | 133.76 |  | 0.40576 | 138.8 |  | 0.33635 | 127.04 |  | 4.3419e+18 | 133.76 |  | 1.1192e+18 | 127.04 |  | 2.0656e+18 | 133.76 |  | 5.3245e+17
146.8 |  | 23.9397 | 153.36 |  | 5.4777 | 156.96 |  | 3912.39 | 163.6 |  | 5.8835 | 166.32 |  | 25.5642 | 173.28 |  | 3.2461 | 180 |  | 0.40576 | 185.04 |  | 0.67271 | 173.28 |  | 3.3082e+18 | 180 |  | 1.1192e+18 | 173.28 |  | 1.5738e+18 | 180 |  | 5.3245e+17
193.04 |  | 19.2735 | 199.6 |  | 5.8835 | 203.2 |  | 4488.68 | 209.84 |  | 9.5353 | 212.56 |  | 18.6813 | 219.52 |  | 2.4345 | 226.24 |  | 0.20288 | 231.28 |  | 0 | 219.52 |  | 2.4811e+18 | 226.24 |  | 5.5959e+17 | 219.52 |  | 1.1804e+18 | 226.24 |  | 2.6622e+17
239.28 |  | 20.085 | 245.84 |  | 6.2892 | 249.44 |  | 6063.46 | 256.08 |  | 9.941 | 258.8 |  | 26.5471 | 265.76 |  | 5.4777 | 272.48 |  | 1.4202 | 277.52 |  | 0 | 265.76 |  | 5.5825e+18 | 272.48 |  | 3.9171e+18 | 265.76 |  | 2.6558e+18 | 272.48 |  | 1.8636e+18
285.52 |  | 19.0706 | 292.08 |  | 8.7238 | 295.68 |  | 7739.1 | 302.32 |  | 12.7814 | 305.04 |  | 52.111 | 312 |  | 8.318 | 318.72 |  | 0.60864 | 323.76 |  | 1.3454 | 312 |  | 8.4771e+18 | 318.72 |  | 1.6788e+18 | 312 |  | 4.0329e+18 | 318.72 |  | 7.9868e+17
331.76 |  | 32.0548 | 338.32 |  | 14.2015 | 341.92 |  | 10445.4 | 348.56 |  | 19.8821 | 351.28 |  | 58.0105 | 358.24 |  | 8.9267 | 364.96 |  | 0.81151 | 370 |  | 1.6818 | 358.24 |  | 9.0974e+18 | 364.96 |  | 2.2384e+18 | 358.24 |  | 4.328e+18 | 364.96 |  | 1.0649e+18
378 |  | 45.8506 | 384.56 |  | 19.6792 | 388.16 |  | 14036.8 | 394.8 |  | 26.9828 | 397.52 |  | 102.256 | 404.48 |  | 15.2159 | 411.2 |  | 2.2317 | 416.24 |  | 2.0181 | 404.48 |  | 1.5507e+19 | 411.2 |  | 6.1555e+18 | 404.48 |  | 7.3773e+18 | 411.2 |  | 2.9285e+18
424.24 |  | 64.9212 | 430.8 |  | 33.2721 | 434.4 |  | 20127.1 | 441.04 |  | 38.3441 | 443.76 |  | 168.133 | 450.72 |  | 27.9972 | 457.44 |  | 4.0576 | 462.48 |  | 4.3726 | 450.72 |  | 2.8533e+19 | 457.44 |  | 1.1192e+19 | 450.72 |  | 1.3574e+19 | 457.44 |  | 5.3245e+18
470.48 |  | 89.2668 | 477.04 |  | 49.7053 | 480.64 |  | 26816.9 | 487.28 |  | 51.7341 | 490 |  | 319.552 | 496.96 |  | 46.6621 | 503.68 |  | 7.9123 | 508.72 |  | 3.6999 | 496.96 |  | 4.7555e+19 | 503.68 |  | 2.1824e+19 | 496.96 |  | 2.2624e+19 | 503.68 |  | 1.0383e+19
516.72 |  | 133.9 | 523.28 |  | 69.5874 | 526.88 |  | 36294.5 | 533.52 |  | 64.3126 | 536.24 |  | 521.117 | 543.2 |  | 68.1673 | 549.92 |  | 7.7094 | 554.96 |  | 8.7452 | 543.2 |  | 6.9471e+19 | 549.92 |  | 2.1264e+19 | 543.2 |  | 3.305e+19 | 549.92 |  | 1.0117e+19
562.96 |  | 144.653 | 569.52 |  | 88.0495 | 573.12 |  | 46812.5 | 579.76 |  | 81.7602 | 582.48 |  | 860.348 | 589.44 |  | 95.3532 | 596.16 |  | 13.39 | 601.2 |  | 12.7814 | 589.44 |  | 9.7177e+19 | 596.16 |  | 3.6933e+19 | 589.44 |  | 4.6231e+19 | 596.16 |  | 1.7571e+19
609.2 |  | 144.856 | 615.76 |  | 86.8322 | 619.36 |  | 55681.1 | 626 |  | 105.497 | 628.72 |  | 1264.49 | 635.68 |  | 125.785 | 642.4 |  | 22.9253 | 647.44 |  | 11.436 | 635.68 |  | 1.2819e+20 | 642.4 |  | 6.3234e+19 | 635.68 |  | 6.0986e+19 | 642.4 |  | 3.0083e+19
655.44 |  | 142.624 | 662 |  | 90.484 | 665.6 |  | 66742.9 | 672.24 |  | 122.336 | 674.96 |  | 1508.34 | 681.92 |  | 137.755 | 688.64 |  | 28.2001 | 693.68 |  | 11.0997 | 681.92 |  | 1.4039e+20 | 688.64 |  | 7.7783e+19 | 681.92 |  | 6.679e+19 | 688.64 |  | 3.7005e+19
701.68 |  | 117.264 | 708.24 |  | 73.645 | 711.84 |  | 73138.3 | 718.48 |  | 132.886 | 721.2 |  | 1778.76 | 728.16 |  | 154.594 | 734.88 |  | 33.8807 | 739.92 |  | 15.1359 | 728.16 |  | 1.5755e+20 | 734.88 |  | 9.3452e+19 | 728.16 |  | 7.4954e+19 | 734.88 |  | 4.446e+19
747.92 |  | 109.149 | 754.48 |  | 78.1084 | 758.08 |  | 85649.3 | 764.72 |  | 150.131 | 767.44 |  | 2096.39 | 774.4 |  | 155.203 | 781.12 |  | 39.5613 | 786.16 |  | 15.4723 | 774.4 |  | 1.5817e+20 | 781.12 |  | 1.0912e+20 | 774.4 |  | 7.5249e+19 | 781.12 |  | 5.1914e+19
794.16 |  | 91.4984 | 800.72 |  | 65.1241 | 804.32 |  | 92525.5 | 810.96 |  | 165.144 | 813.68 |  | 2261.6 | 820.64 |  | 165.955 | 827.36 |  | 39.1556 | 832.4 |  | 11.7724 | 820.64 |  | 1.6913e+20 | 827.36 |  | 1.08e+20 | 820.64 |  | 8.0462e+19 | 827.36 |  | 5.1382e+19
840.4 |  | 83.7891 | 846.96 |  | 61.0665 | 850.56 |  | 105656 | 857.2 |  | 195.982 | 859.92 |  | 2593.01 | 866.88 |  | 161.492 | 873.6 |  | 43.2131 | 878.64 |  | 18.1631 | 866.88 |  | 1.6458e+20 | 873.6 |  | 1.1919e+20 | 866.88 |  | 7.8298e+19 | 873.6 |  | 5.6706e+19
886.64 |  | 66.7471 | 893.2 |  | 41.5901 | 896.8 |  | 112667 | 903.44 |  | 208.763 | 906.16 |  | 2858.54 | 913.12 |  | 145.87 | 919.84 |  | 47.8794 | 924.88 |  | 8.4088 | 913.12 |  | 1.4866e+20 | 919.84 |  | 1.3206e+20 | 913.12 |  | 7.0724e+19 | 919.84 |  | 6.2829e+19
932.88 |  | 48.488 | 939.44 |  | 30.026 | 943.04 |  | 127079 | 949.68 |  | 245.891 | 952.4 |  | 3076.85 | 959.36 |  | 145.87 | 966.08 |  | 43.6189 | 971.12 |  | 7.7361 | 959.36 |  | 1.4866e+20 | 966.08 |  | 1.2031e+20 | 959.36 |  | 7.0724e+19 | 966.08 |  | 5.7238e+19
979.12 |  | 30.6347 | 985.68 |  | 16.2303 | 989.28 |  | 129946 | 995.92 |  | 244.877 | 998.64 |  | 2911.65 | 1005.6 |  | 90.4841 | 1012.32 |  | 31.4462 | 1017.36 |  | 2.6908 | 1005.6 |  | 9.2215e+19 | 1012.32 |  | 8.6737e+19 | 1005.6 |  | 4.3871e+19 | 1012.32 |  | 4.1265e+19
1025.36 |  | 14.2015 | 1031.92 |  | 11.9698 | 1035.52 |  | 126916 | 1042.16 |  | 221.139 | 1044.88 |  | 2678.58 | 1051.84 |  | 53.3571 | 1058.56 |  | 18.0562 | 1063.6 |  | 0.67271 | 1051.84 |  | 5.4378e+19 | 1058.56 |  | 4.9804e+19 | 1051.84 |  | 2.587e+19 | 1058.56 |  | 2.3694e+19
1071.6 |  | 11.1583 | 1078.16 |  | 8.318 | 1081.76 |  | 119396 | 1088.4 |  | 224.994 | 1091.12 |  | 2644.17 | 1098.08 |  | 25.7656 | 1104.8 |  | 7.5065 | 1109.84 |  | 1.3454 | 1098.08 |  | 2.6258e+19 | 1104.8 |  | 2.0705e+19 | 1098.08 |  | 1.2492e+19 | 1104.8 |  | 9.8503e+18
1117.84 |  | 10.9554 | 1124.4 |  | 6.8979 | 1128 |  | 116742 | 1134.64 |  | 209.778 | 1137.36 |  | 2624.49 | 1144.32 |  | 12.5785 | 1151.04 |  | 3.6518 | 1156.08 |  | 0.33635 | 1144.32 |  | 1.2819e+19 | 1151.04 |  | 1.0073e+19 | 1144.32 |  | 6.0986e+18 | 1151.04 |  | 4.7921e+18
1164.08 |  | 7.1007 | 1170.64 |  | 9.5353 | 1174.24 |  | 112143 | 1180.88 |  | 210.59 | 1183.6 |  | 2389.46 | 1190.56 |  | 4.0576 | 1197.28 |  | 0.81151 | 1202.32 |  | 1.0091 | 1190.56 |  | 4.1352e+18 | 1197.28 |  | 2.2384e+18 | 1190.56 |  | 1.9673e+18 | 1197.28 |  | 1.0649e+18
1210.32 |  | 7.3036 | 1216.88 |  | 6.0864 | 1220.48 |  | 107213 | 1227.12 |  | 192.127 | 1229.84 |  | 2448.48 | 1236.8 |  | 1.4202 | 1243.52 |  | 0.60864 | 1248.56 |  | 0 | 1236.8 |  | 1.4473e+18 | 1243.52 |  | 1.6788e+18 | 1236.8 |  | 6.8855e+17 | 1243.52 |  | 7.9868e+17
1256.56 |  | 8.1151 | 1263.12 |  | 6.2892 | 1266.72 |  | 103724 | 1273.36 |  | 180.157 | 1276.08 |  | 2220.31 | 1283.04 |  | 0.60864 | 1289.76 |  | 0 | 1294.8 |  | 0 | 1283.04 |  | 6.2028e+17 | 1289.76 |  | 0 | 1283.04 |  | 2.9509e+17 | 1289.76 |  | 0
1302.8 |  | 5.8835 | 1309.36 |  | 6.8979 | 1312.96 |  | 103313 | 1319.6 |  | 188.881 | 1322.32 |  | 2210.49 | 1329.28 |  | 0.20288 | 1336 |  | 0.40576 | 1341.04 |  | 0 | 1329.28 |  | 2.0676e+17 | 1336 |  | 1.1192e+18 | 1329.28 |  | 98364100000000000 | 1336 |  | 5.3245e+17
1349.04 |  | 8.1151 | 1355.6 |  | 5.8835 | 1359.2 |  | 97334.6 | 1365.84 |  | 171.231 | 1368.56 |  | 2133.79 | 1375.52 |  | 0.40576 | 1382.24 |  | 0 | 1387.28 |  | 0 | 1375.52 |  | 4.1352e+17 | 1382.24 |  | 0 | 1375.52 |  | 1.9673e+17 | 1382.24 |  | 0
1395.28 |  | 8.1151 | 1401.84 |  | 6.2892 | 1405.44 |  | 96340.8 | 1412.08 |  | 183.404 | 1414.8 |  | 2125.92 | 1421.76 |  | 0.20288 | 1428.48 |  | 0 | 1433.52 |  | 0 | 1421.76 |  | 2.0676e+17 | 1428.48 |  | 0 | 1421.76 |  | 98364100000000000 | 1428.48 |  | 0
1441.52 |  | 6.695 | 1448.08 |  | 7.7094 | 1451.68 |  | 97730.3 | 1458.32 |  | 184.215 | 1461.04 |  | 2157.38 | 1468 |  | 0.20288 | 1474.72 |  | 0 | 1479.76 |  | 0 | 1468 |  | 2.0676e+17 | 1474.72 |  | 0 | 1468 |  | 98364100000000000 | 1474.72 |  | 0
1487.76 |  | 4.2604 | 1494.32 |  | 7.7094 | 1497.92 |  | 97182.8 | 1504.56 |  | 187.055 | 1507.28 |  | 2198.69 | 1514.24 |  | 0 | 1520.96 |  | 0 | 1526 |  | 0 | 1514.24 |  | 0 | 1520.96 |  | 0 | 1514.24 |  | 0 | 1520.96 |  | 0
1534 |  | 4.4633 | 1540.56 |  | 5.072 | 1544.16 |  | 94599.4 | 1550.8 |  | 179.752 | 1553.52 |  | 2318.67 | 1560.48 |  | 0.20288 | 1567.2 |  | 0 | 1572.24 |  | 0 | 1560.48 |  | 2.0676e+17 | 1567.2 |  | 0 | 1560.48 |  | 98364100000000000 | 1567.2 |  | 0
1580.24 |  | 6.0864 | 1586.8 |  | 7.7094 | 1590.4 |  | 94674.2 | 1597.04 |  | 182.389 | 1599.76 |  | 2195.74 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1618.48 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0 | 1606.72 |  | 0 | 1613.44 |  | 0
1626.48 |  | 7.5065 | 1633.04 |  | 5.4777 | 1636.64 |  | 99686.8 | 1643.28 |  | 182.795 | 1646 |  | 2215.41 | 1652.96 |  | 0.20288 | 1659.68 |  | 0 | 1664.72 |  | 0 | 1652.96 |  | 2.0676e+17 | 1659.68 |  | 0 | 1652.96 |  | 98364100000000000 | 1659.68 |  | 0
1672.72 |  | 6.8979 | 1679.28 |  | 3.2461 | 1682.88 |  | 98113.3 | 1689.52 |  | 183.404 | 1692.24 |  | 2160.33 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1710.96 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0 | 1699.2 |  | 0 | 1705.92 |  | 0
1718.96 |  | 5.072 | 1725.52 |  | 4.8691 | 1729.12 |  | 97009.3 | 1735.76 |  | 185.635 | 1738.48 |  | 2026.61 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1757.2 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0 | 1745.44 |  | 0 | 1752.16 |  | 0
1765.2 |  | 7.7094 | 1771.76 |  | 6.0864 | 1775.36 |  | 98275.2 | 1782 |  | 181.172 | 1784.72 |  | 2243.92 | 1791.68 |  | 0.20288 | 1798.4 |  | 0 | 1803.44 |  | 0 | 1791.68 |  | 2.0676e+17 | 1798.4 |  | 0 | 1791.68 |  | 98364100000000000 | 1798.4 |  | 0
*** DATA END ***
*** ANALYSIS POSITION ***
Data file name(s) | 12_As-2_NEG GOOD.ckb
Shuttle_ID |  | 18holes
Stage coordinates
X(um) |  |  | -29295
Y(um) |  |  | -13715
Z(um) |  |  | 1789
Window coordinates
X(um) |  |  | 2204
Y(um) |  |  | -15
Sample coordinates
X(um) |  |  | 2204
Y(um) |  |  | -15
*** ANALYSIS RECIPE ***
File name(s) |  | 5 new mass calib now with eGun.rckb
Creation date |  | 2017-09-14
Creation time |  | 14:34
*** MAIN ANALYTICAL PARAMETERS (MAP) ***
Data file name(s)
ISF file name
Primary ions
Primary column
Impact energy (eV)
Secondary ion polarity
Secondary ion energy (V)
*** OPTIONAL INSTRUMENTAL PARAMETERS (OIP) ***
Data file name(s)
Prim beam preset
Prim beam intensity @t=0 (nA)
Prim beam intensity @t=end (nA)
Prim beam intensity regulation
Raster size (um)
DTOS
Analyzed area size (um)
EGate rate (%)
Mass resolution
Field aperture (um)
Contrast aperture (um)
Entrance slit (um)
Exit slit (um)
Energy window (eV)
Energy slit offset (eV)
PBMF aperture (um)
LPRIM4 aperture (um)
Rotation
Speed (rpm)
*** ION DETECTOR PARAMETERS ***
Data file name(s)
EM yield (%)
EM HV (V)
EM discriminator threshold (mV)
EM background (c/s)
EM dead time (ns)
EM -> FC threshold (c/s)
FC -> EM threshold (c/s)
*** OTHERS ***
Data file name(s)
Oxygen flooding
Analysis chamber press (mbar)
Electron gun
e-gun HV
e-beam current
*** CALIBRATION PARAMETERS ***
Data File name(s) | 12_As-2_NEG GOOD.ckb
Layer #1 matrix
1st Interface (nm) | 0
2nd Interface (nm) | 1000
Depth calib method | Depth
Mean SR (nm/min) | 33.27
Species |  | 12C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 13C
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 16O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 18O
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 197Au
Conc calib method
RSF (at/cm3)
Dose (at/cm2)
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal | 28Si
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As_comb
Conc calib method | RSF
RSF (at/cm3) | 2.559e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal
Species |  | 75As 28Si_comb
Conc calib method | RSF
RSF (at/cm3) | 6.926e+21
Dose (at/cm2) | 1.798E+308
[C] (at/cm3)
Ref signal